## BfArM
Base | Device name | Product group | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
BfArM | SPATIALFRAME.COM SOFTWARE | Medical data processing (software) - Other | "It has come to our attention that on rare occasions, when surgeons plan to use the software to adjust struts, the software may produce an "inverted" frame image. It has come to our attention that on rare occasions, when surgeons use the software to plan the adjustment of the struts, the software produces an "inverted" image of the frame."
BfArM | CARESCAPE PDM | Medical Electronics / Electromedicine - Electrodiagnostics | "If the CARESCAPE PDM with Masimo SET® technology has not been completely turned off for more than 2 years, a frozen SpO2 saturation reading may occur that is inaccurate and no longer changes according to the patient's clinical condition. If this situation were to occur, it could result in missed SpO2 alarms and delayed diagnosis and treatment of hypoxia or hyperoxia. "
BfArM | SCENARIA | Radiological Technology - Computer Tomographs | "We have confirmed that the coronary analysis software V1.10, option software for the x-ray CT system SCENARIA can cause a phenomenon in which the ROI settings in the CT value measurement function displays incorrect values of area. This may lead to a misdiagnosis. But the level of harm is negligible; therefore re-examination of patients who have been examined with this function is not necessary. Furthermore, there was no health related hazard reported due to this problem. In order to prevent this phenomenon, modified software, correcting the cause of the problem, will be available soon."
BfArM | ABL700 Series and ABL800FLEX | in vitro diagnostics | "The problem is that calibration errors are not reflected in the Parameter Bar and on the patient results when the calibration number (Cal #) is in the following ranges: 32768 - 65535 98303 - 131070 The overall Analyzer Status (“traffic light”) will, however, correctly reflect the status of the analyzer. Potentially, the software problem may cause the operator to perform measurements, although the analyzer requires intervention, and thus to report questionable results."
BfArM | PAD 300 PAD 300P | Medical Electronics / Electromedicine - Electrotherapy (without physiotherapy) | "Existing Pr oblem points Certain PAD 300 / PAD 300P devices may have the following problems that may affect your ability to provide treatment to a patient in the event of sudden cardiac arrest (SCA) with these devices. patient. Point 1 (Problems with switching on and off on and off): The device could switch itself on without the user having pressed any button. If this is the case, the device will emit the usual sequence of audible signals. If the device does not recognize that the acoustic signals are being followed (e.g. in the form of a patient being connected to the electrodes so that the device is used to retrieve the patient's ECG and that the corresponding actions are carried out in their usual sequence), the device will automatically switch itself off again after ten minutes to save power. This process can be either repeated or intermittent. If this condition remains undetected for a longer period of time, the battery will discharge completely after some time. In the worst case scenario, the battery could discharge to such an extent within just one week that the battery level falls below the minimum capacity required to use the device efficiently in an emergency. In such circumstances, the device can be used again during a subsequent treatment, provided that sufficient power is supplied to the defibrillator. The devices that could potentially be affected by this automatic on/off problem were manufactured between August 2004 and December 2010 and have a guaranteed life of up to three years. a guaranteed life life of up to seven years. H017-101-301-0 Point 2 ( Problem with the battery management software): Certain PAD 300 / PAD 300P devices equipped with older versions of the battery management software may misinterpret a temporary drop in battery voltage and mistake it for an indication of a low battery. This problem then leads to the device displaying the low battery message too early and switching itself off, even though there is still sufficient battery capacity. This can occur when the device is performing its weekly self-test, when the device is switched on, when the device is being prepared to send a power surge, or even after the device has been switched off. If the low battery warning message is triggered due to this problem, the device will still have enough power to run for several more hours. However, it is also possible that the device will only be able to emit a single power surge. However, it is also possible that the device can only emit a single power surge before it switches itself off. If this were to happen, there could be a delay in the delivery of any subsequent power surges. n addition, this power surge would then only be emitted with the lowest current of 150 joules. once, it is more likely that it will notice this again. This can occur if the integrated PAD-PAK has already been completely discharged once due to the normal discharge of the battery or if the integrated PAD-PAK has already been completely discharged once. the integrated PAD-PAK is a battery with a capacity of 800 mAh. This software version was distributed until the end of December 2010. Devices equipped with software version 1.4.2 / 3 .2.0 or a higher version are not affected by this problem."
BfArM | V60 Plus | surgical equipment and anesthesia - Anesthesia and medical gas supply | Stated in "gloves": "Serious or even fatal injuries may occur if users do not respond immediately to device alarms triggered by software version 2.20. version 2.20."
BfArM | Reanibex 300-500 | Medical Electronics / Electromedicine - Electrotherapy (without physiotherapy) | "The REANIBEX 300 and REANIBEX 500 defibrillator models may indicate a sufficient charge for defibrillation between 12 and 50%, but the defibrillator still shuts off suddenly. The defibrillator suddenly switches off. Following an investigation, it has been concluded that the defibrillator's software does not correctly measure consumption when the device is switched off. It can happen that the device does not have sufficient charge for a defibrillation."
BfArM | ADAMS A1C Calibrator 80 | In-vitro diagnostics - Devices / Products for clinical chemistry | "lt was found that "ADAMS A1c Calibrator 80 (INT-IVD Specification)" have an error in the IFCC value described on the Standard Value List and on the barcode Iabei. The subject lots need to be replaced. We appreciate your cooperation."
BfArM | Vivid iq, Vivit T8, Vivid T9 | Ultrasound Technology - US Diagnostic Devices | "GE Healthcare has become aware that certain VIVID ultrasound systems may shut down unexpectedly when images are manually or automatically saved to PACS. This could occur if non-ASCII characters such as "Ø" or "Æ" are included in the descriptions or accession number of the scheduled procedure in the DICOM Modality Worklist. d. This situation is unlikely. h. Should this situation occur and the system shuts down during an interventional procedure, this could result in a loss of imaging during the procedure. No injuries have been reported in connection with this problem."
BfArM | Vivid S60N, S70N, E80, E90, E95 | Ultrasound Technology - US Diagnostic Devices | "GE Healthcare has become aware that certain VIVID ultrasound systems may shut down unexpectedly when images are manually or automatically saved to PACS. This could occur if non-ASCII characters such as "Ø" or "Æ" are included in the descriptions or accession number of the scheduled procedure in the DICOM Modality Worklist. This situation is unlikely. Should this situation occur and the system shuts down during an interventional procedure, this could result in a loss of imaging during the procedure. No injuries have been reported in connection with this problem. "
BfArM | Vscan Extend | Ultrasound Technology - US Diagnostic Devices | "GE Healthcare has become aware of an issue where the Vscan Extend will not turn on due to a depleted internal battery. This problem can occur if the VSCAN is not turned off before the external battery is removed or if the device is connected to a PC but not connected to power. In such situations, the Vscan Extend cannot be used until it is connected to power and has been charging for approximately 10 minutes. This could lead to a delay in treatment."
BfArM | Giraffe Inkubatoren and Giraffe OmniBeds | Physical Therapy - Climate Therapy | "GE Healthcare has determined that certain Giraffe Incubators and Giraffe OmniBeds manufactured prior to 2012 may deliver a different oxygen level to the environment than that displayed by the device when a certain optional component, called the Servo Oxigen Module, is installed in the device. The Servo Oxigen module automatically regulates the oxygen delivered internally by the device. Standard pulse oximetry monitoring detects changes in oxygen levels in the patient's blood, so this problem should only result in transient hypoxia or hyperoxia. No injuries have been reported in connection with this problem. "
BfArM | CARESCAPE R860, ENGSTRÖM CARESTATION, ENGSTRÖM PRO | surgical equipment and anesthesia - Anesthesia and medical gas supply | "The connector housing on the neonatal flow sensor cable may be missing, which could expose the wires in the cable. If the exposed wires break, primary ventilation parameters may be lost (flow and tidal volume). This problem can cause indirect but reversible changes to the patient's condition, as these parameters are used alongside other information to select the correct ventilator settings for the patient. No injuries have been reported in connection with this problem."
BfArM | HemochronTM Signature Elite | In vitro diagnostics - Products of hematology, histology and cytology | "This letter is to inform you and your facility that Accriva Diagnostics has determined that some Hemochron Signature Elite analyzers are not maintaining 37°C ±1.0°C degrees in some operating environments as described in the Operator's Manual (P/N HX1101). Internal testing has shown no clinically significant effect on the analytical performance of any parameter when the temperature exceeds 37 °C ±1.0 °C. For this reason, this notice is intended to inform you of this limitation with respect to current temperature requirements. Software version 2.4 is designed to ensure that the Hemochron Signature Elite analyzer maintains a temperature of 37 °C ±1.0 °C throughout the measurement process. Your analyzers will be upgraded to software version 2.4 during the next maintenance or repair. as part of the next maintenance or repair."
BfArM | Centricity Universal Viewer Zero Footprint Client (ZFP) | Medical data processing (software) - Data processing systems for image processing | "GE Healthcare has become aware of an issue where distance and area measurements may be displayed inaccurately on images that have been scaled down from their original resolution."
BfArM | Merge Eye Station und Eye Care PACS | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "This message was issued due to a potential issue that prevents image capture or import. This issue was reported by a customer whose configuration allows sharing a database from a deprecated version and a new WSM (WinStation Manager) version. Another symptom involved software preventing the merge capture session from opening. Use of this product may delay patient treatment and/or diagnosis"
BfArM | Centricity PACS | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "The Event Notification Manager (ENM) feature for certain Centricity PACS products does not process notifications for study changes made to post-verified studies. If the study/series split feature is used to correct mismatched patient images within a study/series, these changes will not be transmitted to the Centricity Enterprise Archive or to Enterprise Archive for specific Centricity PACS products. Viewing Centricity Enterprise Archive or Enterprise Archive directly with a diagnostic viewer and in rare cases where the issue is not recognized by the physician may result in studies being displayed with incorrect patient images or patient information. Note: The series and/or study changes are properly updated in the Centricity PACS database. This issue does not affect images viewed through a diagnostic viewer that is directly connected to the Centricity PACS Foundation. No injuries have been reported related to this issue."
BfArM | MAC VU360 | Medical Electronics / Electromedicine - Electrodiagnostics | "The MACTM VU360 system may intermittently display an incorrect patient ID or visit number on the screen after scanning the patient's barcode. This may result in the ECG report being assigned to the wrong patient. pt.1 pt.2: The MACTM VU360 may display incorrect patient demographic information on the patient banner. This may result in the ECG report being assigned to the wrong patient. For points 1 and 2: If the patient information is not matched to the correct ECG, it is possible that the patient will receive an incorrect diagnosis and/or unnecessary treatment and/or that there will be a delay in finding the correct diagnosis and performing it the necessary treatment can come. No injuries were reported as a result of these issues."
BfArM | 1 Patient Information Center iX 2 Erweit. Patient Information Center iX 3 PIC iX Essentials 4 PIC iX Essentials Erweiterung | Medical Electronics / Electromedicine - Electrodiagnostics | "Philips has received reports that PIC iX 4.x Surveillance or Overview products may crash or freeze with a blank screen. This is caused by an interaction between the software process of the Windows operating system and the Intel graphics driver on the PC host on which the PIC iX 4.x application software is running. If this issue occurs, the PIC iX 4.x Surveillance or Overview products will require a manual hardware reboot to restart and resume central patient monitoring. It is important to note that if the problem occurs, any connected bed monitors or telemetry devices will detect that the connection to the PIC iX has been lost and will display a fault message or technical alarm “No Central Monitoring”."
BfArM | CARESCAPE ONE | Medical Electronics / Electromedicine - Electrodiagnostics | "If a patient has a ventricular fibrillation arrhythmia at the time of docking CARESCAPE ONE to a CARESCAPE B450/B650/B850 host monitor and the host monitor is in the CARESCAPE ONE learning process, the host monitor may not respond to ventricular fibrillation. indicate arrhythmia. In rare circumstances, this problem could lead to delayed clinical assessment of V-fib arrhythmia. No incidents or injuries have been reported as a result of this issue. A NOTICE: 1) This issue does not affect the detection of the V-Fib arrhythmia alarm before docking the CARESCAPE ONE to the host monitor or after the CARESCAPE ONE is docked to the host monitor and the learning message is no longer displayed. 2) The problem does not affect the ECG waveforms or other alarms, e.g. B. Pulse rate high/low or ECG heart rate high/low alarms derived from the SpO2 sensor."
BfArM | Thermo Scientific lndiko, lndiko Plus, Gallery, Gallery Plus and Gallery Plus Beermaster analyzers | In-vitro diagnostics - Devices / Products for clinical chemistry | "During the internal test runs a bug in the Software with potential risk of mis- match of dispensed reagent was identified. This causes a potential risk of wrongly reported results for the lndiko and Gallery basic models. ln the Plus models there is no risk for wrongly reported results."
BfArM | CARESCAPE PDM | Medical Electronics / Electromedicine - Electrodiagnostics | "The CARESCAPE Patient Data Module (PDM) is used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro. When pacemaker detection is turned on on the monitor and an automated external defibrillator (AED) is used for defibrillation, the low frequency signals from the AED cause the PDM to produce a false baseline ECG waveform and display the asystole alarm on the monitor. This problem may result in delayed clinical evaluation of an ECG arrhythmia event. No incidents or damage have been reported as a result of this issue. A NOTICE: • This problem only occurs when an AED is used for defibrillation. If a defibrillator If used in hospital mode (not AED mode), this problem will not occur. • This problem affects the efficiency of the AED in diagnosing the patient's rhythm or administering the required defibrillation therapy. • This issue does not affect CARESCAPE PDM pulse oximetry monitoring."
BfArM | PACS-IW | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "Your site has received the UV-IW Image Acquisition and Synchronization (IAS) Tool as part of GEHC Ref# 85456. When operating correctly, the IAS tool identifies incomplete studies, temporarily hides them, and makes them available for interpretation without unnecessary delay when the studies are complete. GE Healthcare has discovered that after installing the IAS tool, some studies may remain hidden without user notification. Access to these hidden studies requires assistance from GE Healthcare, which may result in a delay in reading the study. No injuries have been reported related to this issue."
BfArM | Centricity High Acuity Critical Care (CHA CC) and Centricity High Acuity Anesthesia (CHA A) | Medical Electronics / Electromedicine - Electrodiagnostics | "In rare cases, due to a software error, the CHA application may display an overdue task notification for a scheduled task (e.g., as shown in Figure 1) that has already been completed and documented in the system. This issue can only occur if the Order Module feature is enabled in the CHA system and users have created orders with scheduled tasks. The overdue task notification will appear 30 seconds after the task was documented and cannot be deleted or removed. In rare cases, if this duplication is not noticed, it could potentially lead to inappropriate treatment (e.g. administration of an additional dose of medication or duplicated task). No injuries have been reported related to this issue."
BfArM | 1.5T SIGNA HDxt SIGNAWorks Edition (2019 Upgrade für 1.5T SIGNA HDxt-Systeme) SIGNA Architect SIGNA Architect SIGNA Pioneer SIGNA Premier | Electro-Magnetic Fields - Devices for MR-Tomography | "It was discovered that due to a possible issue with the installation workflow, the MR system date may be set incorrectly. The system's date and time settings are used to populate the DICOM header information on the images. This could result in inaccurate dates on the images. No injuries were reported."
BfArM | OrthoScan HD 1000-0004 Windows XP OS OrthoScan HD FD 1000-0004 Windows XP OS OrthoScan Mobile DI 1000-0005 Windows XP OS OrthoScan 1000-0004-FD Windows XP OS | Radiological Technology - Mobile X-ray diagnostic facilities | "OrthoScan has determined that there is a potential hazard that requires the user not to select the digital zoom feature when the fluoroscopic X-ray beam is active. When selected while the X-ray beam is activated, the digital zoom function prevents the system from displaying the full active X-ray field on the monitor, with the portion of the active image receptor field of view not displayed on the monitor exceeding 3% of the SID (Source -to-Image-Distance – source-to-image distance), which violates DIN EN 60601-2-54 (IEC 60601-2-54)."
BfArM | Mavidon | "Mavidon Issues Voluntary Worldwide Recall of All manufactured products including LemonPrep® Tubes and Single use cups, PediaPrep® Tubes and Single use cups, Wave Prep® Tubes and Single use cups, Cardio Prep Single use cups due to possible Burkholderia cepacia contamination" | "We were notified on December 19, 2019 that samples of 114gm tubes of Lemon Prep, collected during a Food and Drug Administration inspection that occurred at our facility on October 15, 2019 were tested and found to be contaminated with Burkholderia cepacia. Out of an abundance of caution, we are recalling all products manufactured at our facility."
BfArM | CARESCAPE ONE Monitor | Medical Electronics / Electromedicine - Electrodiagnostics | "The CARESCAPE ONE monitor may lose ECG parameters and waveforms when used with an automated external defibrillator (AED) or an AED-capable defibrillator in AED mode. Should this situation occur, the CARESCAPE ONE will display the technical alarms “ECG Malfunction” and “Check Device”. This problem can result in delayed investigation of a life-threatening ECG arrhythmia event. No injuries have been reported related to this issue. NOTE: This issue does not affect the effectiveness of the AED in diagnosing the patient's ECG rhythm or administering the required defibrillation therapy. This issue does not affect the SpO2 or CO2 parameters on the CARESCAPE ONE."
BfArM | Centricity Universal Viewer | Medical data processing (software) - Data processing systems for image processing | "There is a software anomaly in the study management feature of the Centricity Universal Viewer that prevents study changes from being transferred to the Centricity Enterprise Archive (EA) or another Vendor Neutral Archive (VNA). This means that studies with incorrect patient images or patient information can be displayed if these studies are accessed directly from the archive. This issue occurs when 1) the Centricity Universal Viewer is configured to synchronize between the Centricity PACS backend and either the Centricity EA or another VNA AND 2) when a diagnostic viewer (e.g. Centricity Universal Viewer Zero Footprint or a third-party viewer) connects directly to the archive to view images. Both 1 and 2 are required for the problem to occur. NOTE: The series and/or study changes are properly updated with the Centricity PACS database in the Centricity Universal Viewer. This issue does not affect images viewed by Centricity Universal Viewer, Centricity Universal Viewer Zero Footprint, or the RA1000 workstation that is directly connected to the Centricity PACS Foundation. No injuries have been reported related to this issue."
BfArM | ClearCanvas RIS/PACS | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "This recall was triggered due to a software bug in the product that occurs when used with non-DICOM compatible JPEG 2000 compressed images. The software incorrectly uses the DICOM bit depth (i.e., the Bits Stored Tag) to decompress the compressed pixel data stream for display, rather than the bit depth encoded in the compressed pixel data stream. In non-DICOM compatible images where the bit depths of DICOM and the compressed pixel data stream are different, the software outputs an image with some loss of accuracy of the decompressed pixel data. When JPEG 2000 compressed images with the characteristics described above are displayed with the affected product, the images initially appear too dark (or too light), or in many cases completely black (or white). The ability to detect subtle differences in contrast between adjacent pixels in the image may be impaired. If the user attempts to adjust the window and plane to visualize the anatomy, they will find that there is a higher sensitivity than usual and it is more difficult to obtain an image of acceptable quality or an image with subtle differences in contrast."
BfArM | Centricity Universal Viewer mit PACS-IW Foundation 6.0 und Centricity PACS-IW mit Universal Viewer Version 5.0 | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "It is possible that one or more images or image series are missing from exams without a warning being displayed in the viewer. There are two situations that can cause errors in image capture so that some images are not displayed: 1. Error in image compression or mapping. 2. Images that have more than 128 characters in the Image Comment attribute (0020 4000) contained within the DICOM header. There have been no reported personal injuries associated with this issue."
BfArM | AngioDynamics Solero-Mikrowellen-Gewebeablationssystem (MTA): Generator und Applikator | Medical Electronics / Electromedicine - Electrotherapy (without physiotherapy) | "For your information: Risk Management Department AngioDynamics, Inc., the manufacturer of the Solero Microwave Tissue Ablation System (MTA), the Solero Generator, and the Solero Applicators, is conducting a consumer Field Safety Corrective Action (FSCA) of which certain serial numbers and lots/lots/ Lots of these devices are affected. The FSCA requires replacement of the instruction manual/instructions for use provided with the device to reduce the risk of potential harm to patients due to delay in performing a procedure. This Field Safety Notification (FSN) will be sent to all recipients of the Solero generator and applicator and will include the revised Operator's Manual/Instructions for Use. The instructions and warnings have been made clearer to ensure that the Solero system is functioning properly before starting an intervention. The current Solero Microwave Tissue Ablation Generator Manual (“System Manual”) and the Solero Microwave Tissue Ablation System Applicator Instructions for Use (“Applicator Instructions for Use”) contain the following indications: “The Solero Microwave Tissue Ablation System (MTA) and accessories are intended for soft tissue ablation* in open, laparoscopic and percutaneous procedures. The Solero MTA System is not intended for use in cardiac testing.” The Solero System Manual and Applicator Instructions for Use both contain the following warning statement: WARNING: “Do not place the applicator in the patient until it is connected, successfully bled, and the generator status bar displays Ready.” The above warning is intended to ensure that the Solero System starts properly and is fully functional prior to placement in the patient. However, based on user reports, we have determined that this warning could clarify the use of the applicators in open procedures. The alert has been revised as follows to ensure the system is functioning properly prior to surgical preparation of patients. This reduces potential patient harm (e.g., from prolonged exposure to the anesthetic and/or premature initiation of surgery) due to a delay in therapeutic care. (Revised) WARNING: “Do not initiate the procedure or anesthesia until the applicator is connected, primed and the generator status bar displays Ready.” Please review the System Manual and Applicator Instructions for Use included in this package and read the added statement in the Warnings section. There is no need to decommission or return any Solero generators/applicators as part of this corrective action. Page 1 of 3 In addition to these changes to the User Manual and Instructions for Use, AngioDynamics will also install a software update that will reduce the occurrence of Error 0001 that may occur during the initial boot process. This improvement will be performed the next time the Solero Microwave Tissue Ablation System Generator is serviced. AngioDynamics has confirmed that the generators and applicators affected by this corrective action have been shipped to end users worldwide. AngioDynamics began shipping the products affected by this corrective action on Tuesday, February 28, 2017. Our records indicate that your healthcare facility has received one or more of the devices affected by this corrective action. For complete details of the affected product that was delivered directly to your organization, please refer to the Response Confirmation Control Document included with this FSN."
BfArM | Standard C Cem standard, Size 8 Standard C Cem lateralized, Size 8 | Non-active implants - Bone surgery | "The Original Equipment Manufacturer (OEM) informed us that he has not renewed the CE certification for the stems with size 8. We decided to recall these stems as a precautionary measure as we have no evidence that they cornply with current standards."
BfArM | Centricity Universal Viewer Zero Footprint Client (ZFP), Centricity PACS RA1000 Workstation (RA1000), Centricity Radiology RA600 (RA600), Centricity Cardiology CA1000 (CA1000) und Centricity Enterprise Web (CWeb) | Medical data processing (software) - Data processing systems for image processing | pt.1: "Distance and area measurements can show inaccurate measurements, if these are carried out with enlarged images. Even a printout on film/paper in the original size of the images reflects these inaccurate measurements. Images exported from RA600 or CA1000 to storage media (e.g. CD) will also reflect these inaccurate measurements. The measured value is always overestimated (the measured quantity is larger than the true quantity). In the unlikely event that this problem goes unrecognized, it could potentially lead to inappropriate medical treatment. No injuries have been reported related to this issue."; pt. 2: "Specifically for NDT measurements, distance and area measurements can show inaccurate readings when performed on lossy images that are downscaled from their original resolution. In the unlikely event that this problem goes unrecognized, it could potentially lead to inappropriate medical treatment. No injuries have been reported related to this issue."
BfArM | Extreme H2O 59% | Ophthalmological technology - Contact lenses | "It was discovered that some lenses were inadvertently mislabeled in Lot #0114511565. The lens power stated on the label is +3.75 D, while some of the lenses in the blister packs are actually -2.00 D."
BfArM | ApexPro Telemetrieserver | Medical Electronics / Electromedicine - Electrodiagnostics | "GE Healthcare has become aware of six potential safety issues that may occur when the ApexPro Telemetry System experiences any of the following conditions: NO TELEM, ECG LEADS FAIL, system clock changes, or System restarts. In order for the problem in question to occur, a specific sequence of conditions is required, which are described below. Problems 1-2 can affect ECG monitoring. Issues 3-6 may affect SpO2 monitoring."; https://www.bfarm.de/SharedDocs/Kundeninfos/DE/10/2019/15041-19_kundeninfo_de.pdf?__blob=publicationFile
BfArM | Centricity PACS-IW V3.7.x, Centricity PACS-IW with Universal Viewer V5.0 and Centricity Universal Viewer V6.0 with PACS-IW Foundation | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "GE HealthCare has become aware of an issue where images from two different patients may be included in a single study when used in Centricity PACS-IW V3.7.x, Centricity PACS-IW with Universal Viewer V5.0, and Centricity Universal Viewer V6 .0 were saved with PACS-IW Foundation. This problem could lead to an unfavorable outcome for the patient."
BfArM | CentricityTM Enterprise Archive V2.X, V3.X, V4.X and Enterprise Archive Version 8 | Radiological Technology - Additional devices and accessories for d. X-ray diagnostics/nuclear medicine | "GE HealthCare has become aware of an issue where images from two different patients may be included in a single study when stored in Centricity Enterprise Archive (EA) and/or Enterprise Archive (EA). This problem could contribute to an unfavorable outcome for the patient. This situation can occur during the following workflow: 1. The capture device (e.g., a DX scanner) receives the information about the scheduled procedures (also referred to as orders), including patient demographics, accession number, procedure code, and study instance UID from a DICOM modality worklist , provided by a radiology information system (RIS), cardiovascular information system (CVIS), EMR, or interface module. 2. The healthcare professional incorrectly selects Patient A on the capture device when performing an examination on Patient B. 3. The healthcare professional recognizes this error and attempts to correct the problem by correcting the study patient demographics to Patient B, but this does not change the study instance UID. 4. When Patient A arrives, the healthcare professional selects the order for Patient A on the capture device and saves the study in EA. Due to the same study instance UID, studies for Patients A and B are merged (i.e., images from both Patient A and Patient B are included), taking Patient A's demographics into account."
BfArM | LigaSureTM (Short) Instruments | Optics / Precision Mechanics | "Tyco / Valleylab has identified a possible inadequacy with LigaSureTM Atlas instruments. Certain lots of LigaSure AtlasTM instruments may not be recognized by the LigaSureTM system. The plug indicator then does not change from red to green. Additionally, there may be a slight risk that an incomplete seal may occur. An incomplete seal can pose a risk to the patient."
BfArM | Vario-Cup Prosthesis - seif-eentering Prosthesis Heads LCU Cementless Hip Prosthesis System Gemini SL tibia components | Non-active implants - Bone surgery | "lt has been noticed that the packaging system for the listed Iot and serial numbers does not meet the specifications. A human error led to incorrectly applied sealed seams. The sterility of the products is maintained as the sterile barrier is provided by at least one pouch."
BfArM | LeMaitre Over-The-Wire-Embolektomiekatheter, 5F Plus | Injection / Infusion / Transfusion / Dialysis - Catheter | "Complaint data from 2020 indicates an increased trend for ligature slippage or tip detachment."
BfArM | HAMILTON-G5/HAMILTON-S1 | surgical equipment and anesthesia - Anesthesia and medical gas supply | "Studies have shown that after using the suction function, including disconnecting the patient, suctioning and reconnecting the patient, the preset ventilation pattern may not continue as expected. Under certain conditions, a ventilation mode other than the one set by the user can be applied; this situation can occur regardless of the selected patient group (neonatal, pediatric and adult). This case can occur with the HAMILTON-G5 and HAMILTON-S1 when the ventilation mode changes during preoxygenation. After the patient has been suctioned off and reconnected, the ventilation mode that was active before pre-oxygenation is applied. Under these circumstances, the displayed mode does not match the applied ventilation mode."
BfArM | Standard C Cem standard, Size 9 and 10 Standard C Cem lateralized, Size 9 and 10 | Non-active implants - Bone surgery | "Waldemar Link received an information that the Original Equipment Manufacturer (OEM) of this stems introduced a weight Iimitation of 66 kg for the sizes 9 and 10 of the stems to comply with current standards. The sizes 11 - 18 are not affected of this change. The Waldemar Link GmbH & Co. KG follows the OEM and publishes this change via a field safety notice."
BfArM | Mobilelink lmpactor Handle Stainless Steel, L= 406 mm | Human Medical Instruments - Bone Surgery | "lt is noticed that a Mobilelink impactor handle could not be disassembled from an implanted cup intraoperatively. The characteristics of the thread of the impactor handle do not meet the requirements and caused this problem . Three batches are affected ."
BfArM | Ende-Model Knee Fusion Nail, Femoral component left Size 1, cementable | Non-active implants - Bone surgery | "lt was noticed that one Ende-Model Knee Fusion Nail, Femoral component left Size 1, cementable could intraoperatively not be secured with the assembly screws to the Tibial component. This incident was caused by a manufacturing error which led to a misalignment of the boreholes of the Femoral component. Only one batch is affected."
BfArM | CARESCAPE Telemetrieservers, Apex Telemetrieservers, der CARESCAPE Zentralstation (CSCS) Version 1 und der Central Information Center (CIC)-Systeme | Medical Electronics / Electromedicine - Electrodiagnostics | "Certain versions of the CARESCAPE Telemetry Server, Apex Telemetry Server, CARESCAPE Central Station (CSCS) Version 1, and Central Information Center (CIC) systems have been identified as having vulnerabilities that may affect them when connected to Mission Critical (MC) and/or making the Information Exchange (IX) network vulnerable to cyberattacks. The MC and IX networks are isolated from other hospital networks and traffic. Therefore, for this issue to occur, the unauthorized person would need to gain physical access to the monitoring devices themselves or direct access to the isolated MC or IX networks on-site at the hospital. If an unauthorized person with specialized knowledge gains such access, a combination of a disclosed private key, vulnerable services and components with identified software vulnerabilities could potentially be exploited and combined with further targeted malicious actions to: - Make changes at the operating system level of the device, which, for example, causes the device to B. becomes unusable, and/or - Use services designed to remotely monitor and control devices on the network to access the clinical user interface to make changes to device settings and alarm limits. In such a situation, such cyberattacks could potentially result in reduced monitoring and/or reduction in alerts during active patient monitoring. There have been no such incidents of clinical cyberattacks or injuries reported due to this issue."
BfArM | Tempus Pro (with Triceps 7-Hardware) | Medical Electronics / Electromedicine - Electrodiagnostics | "An unintentional software error was discovered as part of an internal production control for the Tempus Pro. Upon tracing the original identified error, it was determined that this is an intermittent software error that can occur immediately after unplugging the Video Laryngoscope (Part Number: 8403XSB) from the USB port. The software error may impact the Tempus Pro's ability to continue patient monitoring after the video laryngoscope is unplugged. The software error only occurs on Tempus Pro devices with Trizeps 7 hardware. Reviewing our post-market monitoring data, we identified only one complaint related to this issue. There have been no reports of patient harm resulting from this issue."
BfArM | 8240 Mycoplasma genitalium Control Panel (Inactivated Pellet) | In vitro diagnostics - Microbiology products | "While analyzing the long-term stability results for 8240, we noticed some anomalies in the data that triggered a review of the product's performance. The 8240 was designed as a highly diluted control and it was found that our QC assay was unable to measure at the low end of our specification (late Ct). A recent investigation revealed that the above batches were prepared at concentrations that may not be detectable in some M.genitalium assays. For PCR-based assays with a Ct cutoff at or before 33 Ct, false negative or “indeterminate” results may occur with this product. This product is an untested control and users are responsible for ensuring compatibility with their assays. This product is an untested quality control for professional use, i.e. H. it is not intended for a specific assay. The end user can use the control in a laboratory developed test or an assay of their choice. The package insert notes that it is the customer's responsibility to ensure compatibility of the control with the assay or protocol used. Problems may arise with assays with an early Ct cutoff or a LOD below Ct=33."
BfArM | ENROUTE® TRANSKAROTIDALES STENTSYSTEM | Non-active implants - Special implants | "The nose cone (distal tip) may become detached during use of the ENROUTE Stent Delivery System."
BfArM | CoreHip NT1136R, NT1137R, NT1107R | Human Medical Instruments - Bone Surgery | "We internally discovered that it can be possible that the CoreHip Trial Neck Adaptersand CoreHip Primary Calcar Cutting Guide do not engage completely with tlle rasps. tllerefore. a displacement of maximum 1.7mm may occur."
BfArM | Carbon dioxide for laparoscopy | Physical Therapy - Insufflation Technology | "We have noticed that new gas cylinders with a certain valve batch often exhibit leaks at the seat of the rupture disk (which may only occur with a delay), so that in the worst case the gas cylinders can be completely emptied. So the potential error lies in the packaging. These are exclusively certain gas bottles with a nominal volume of 10 liters or a capacity of 7.5 kg. The quality and safety of the medical device (carbon dioxide for laparoscopy) are not affected. If you have already used the product, please note that the medical product itself is of the usual high quality and that there was or is no danger during use for the patient or the staff. We were able to identify all affected gas cylinders and the customers supplied with them in our traceability system. Unfortunately, this affects you, which is why we ask you to return the gas bottles mentioned to us in consultation with our logistics department and to arrange a prompt and free replacement. Detailed information about the recall can be found in the manufacturer's official Field Safety Notice attached to this letter."
BfArM | WELCH ALLYN AED 10 AND MRL JUMPSTART DEFIBRILLATORS | Medical Electronics / Electromedicine | "Dear customer, Our records indicate that you purchased a Welch Allyn AED 10 or MRL JumpStart defibrillator, which is part of our safety corrective action. Depending on the date of manufacture, there is a small possibility that your device may have one or more of the following problems (see Appendix A): • Electric shock energy too low • Electromagnetic noise interference • Unexpected shutdown during use • Blown fuse • Loss of voice prompts • Switch off when cold If you are not the person responsible for this device, please ensure that this letter is forwarded to all employees within your organization who either use or manage this device and should be aware of this security notification."
keywords: "recall, software"

## FDA MAUDE (since 01 01 21)
Base | Device name | Product group | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD CRLSYS KIT | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported on an unknown date in 2023, that the surgeon was programming the autostrut head unit.Motor number 3 would not pass the initial test in the first screen in the software.The surgeon was advised to check the motor clip to make sure the motor was sitting in the strut appropriately, he confirmed it was.He then ran manual mode forward and backwards one time to see if the motor was turning the gear in the strut.It was confirmed the motor was lighting up green and it moved.He then tested again, but the strut did not pass.He replaced the short strut to see if the strut was the issue, but it still was not passing the initial test.The motor was removed and used the blue knob to be sure there was not too much tension on the strut.Confirmed there was not too much tension.The surgeon then used motor number 4 in strut number 3 during the test to see if it could pass, and it did.It was advised at that time to replace the head unit.The new head unit was applied and loading of the program passed and was successful.This report is for one (1) maxframe autostrut(tm)hexapod crlsys kit.This is report 1 of 2 for complaint (b)(4)."
FDA MAUDE | UPGRADE SLEEPWORKS 9.X ACQUISITION SOFTW | SOFTWARE FOR DIAGNOSIS/TREATMENT | "Part lc201-u9 upgrade sleepworks 9.X acquisition software with spike / event and high res video.Issue reported that when marking events on a patient in a review screen the events will also be added to a different patient in the live acquisition window.This could lead to an improper diagnosis."
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD STRUT-MED | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that the patient stated he was sliding down the stairs to get to the main floor and strut number 5 had broken.The patient's mother stated that dontorious bangs the strut on every step on the way down.The mother zip-tied the strut to the ring, and it was broken for a total of two days.The surgeon checked the strut lengths of all six struts, and the program had kept them within 1-2mm of the planned robotic schedule.Strut number 5, even though it had broken two days prior to the strut change, was also within the 1-2mm difference, and the surgeon felt it had no negative impact on patient care.Strut 5 was a medium strut, and it broke through the shoulder bolt knuckle connection.The product is available for return.The scheduled strut change was successful, and no harm to the patient was observed.Strut 5 was a scheduled strut change on this date, and it was replaced with a large auto strut.This also needed a new shoulder bolt because of the strut failure, so they will have both the strut and the shoulder bolt available for return.This report is for one (1) maxframe autostrut(tm)hexapod strut-med this is report 1 of 1 for complaint (b)(4)."
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD STRUT-MED | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported on an unknown date, the patient came into clinic with three broken struts.The patient claimed that he was sitting by space heater when three of his medium autostruts cracked at the base of bolt connection.He is not bearing any weight and this is the second occurrence of this event.Patient claimed to not be tightening the struts himself.There was no patient outcome or consequences.This report is for one (1) maxframe autostrut(tm)hexapod strut-med this is report 3 of 3 for complaint (b)(4)."
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD STRUT-MED | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported on an unknown date in 2023, that the surgeon was programming the autostrut head unit.Motor number 3 would not pass the initial test in the first screen in the software.The surgeon was advised to check the motor clip to make sure the motor
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD STRUT-MED | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported on an unknown date, the patient came into clinic with three broken struts.The patient claimed that he was sitting by space heater when three of his medium autostruts cracked at the base of bolt connection.He is not bearing any weight and this is the second occurrence of this event.Patient claimed to not be tightening the struts himself.There was no patient outcome or consequences.This report is for one (1) maxframe autostrut(tm)hexapod strut-med.This is report 1 of 3 for complaint (b)(4)."
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD CRLSYS KIT | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that on (b)(6) 2023, that autostrut head unit was malfunctioning.Strut #5 malfunction.There was no surgical delay.The surgery was successfully completed.This report is for one (1) maxframe autostrut(tm)hexapod crlsys kit this is report 1 of 2 for complaint (b)(4)."
FDA MAUDE | MAXFRAME AUTOSTRUT(TM)HEXAPOD STRUT-LNG | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that on (b)(6) 2023, that autostrut head unit was malfunctioning.Strut #5 malfunction.There was no surgical delay.The surgery was successfully completed.This report is for one (1) xframe autostrut(tm)hexapod strut-lng.This is report 2 of 2 for complaint (b)(4)."
FDA MAUDE | SMART TSF.COM | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that, after an external fixation procedure with a taylor spatial frame, done on (b)(6) 2022, to correct an equinus deformity, and after a couple of post-op visits, the patient complained of excessive pain, and wanted to have the frame removed.The patient removed a portion of his ring off and was unable to follow the adjustment schedule throughout, the first 40 days scheduled by the doctor were completed in 10 days.The post operative adjustment schedule made with the smart tsf.Com web application was revised by the tsf ally team, and had to be done multiple times due to errors made by the doctor and accidental deletions of the program.During a troubleshooting call made with the surgeon on (b)(6) 2022, it was noticed that the doctor made an error when converting from x-ray mode (program provided by tsf ally team) to manual mode, also strut settings did not match any day of the provided prescription, in some cases, struts were off by over 20mm (struts 3&4).The patient had the taylor spatial frame removed on (b)(6) 2022, no long term harm was reported.It is unknown if additional treatment is necessary, but the patient has refused further surgery."
FDA MAUDE | SMART TSF.COM | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that, at the moment of trying to import an x-ray ((b)(4)) into the smart tsf software, and have the software detect the beacon on the lateral view, the software did not recognized the image.It was cropped the anterior part of the image away ((b)(4)), and it was able to get the software to find the beacon in the smart tsf account.Unfortunately, this image did not work for the physician.Surgeon complained about wasted time (3hr) in clinic, and unwanted radiation exposure to the patient."
FDA MAUDE | SMART TSF.COM | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that, during an internal fixation procedure, the physician was trying to program in the smart tsf.Com a proximal reference case after building a frame with a short proximal ring block.The physician saw 4 dots in 3 arms of the beacon, but the software could not detect the beacon.Physician decided to use the manual mode instead.Surgery was resumed, without any delay, with a back-up device.Patient was not harmed."
FDA MAUDE | SPATIALFRAME.COM SOFTWARE | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that the surgeon wanted to use software for treatment of the taylor spatial frame software.The surgeon navigated to www.Spatialframe.Com at the beginning of a surgery case, because the surgeon wanted to use the chronic mode in the software to ensure he was building an optimized frame.Both the surgeon and the sales representative were unable to login.The incident occurred prior to logging in.The login screen was not available.There was a delay greater than 30 minutes for the adverse event.No injury reported."
FDA MAUDE | SPATIALFRAME.COM SOFTWARE | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that software programming issue led to a patient revision surgery due to a translated fracture.The malfunction reported happen due to mounting parameter error, the original hardware portion of the device was used.However, a new software program was needed."
FDA MAUDE | ORTHOFIX | SOFTWARE FOR DIAGNOSIS/TREATMENT | "Device not recording usage properly.Fda safety report id# (b)(4)."
FDA MAUDE | SPATIALFRAME.COM SOFTWARE | SOFTWARE FOR DIAGNOSIS/TREATMENT | "It was reported that in the taylor spatial frame software, a hardware option was not available for selection.No patient involvement reported."
in other category (Software Algorithm Device to Assist Users In Digital Pathology) there has been no results

## FDA MDR 2021
Base | Device name | Device use | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
FDA MD Recalls | Datascope/Getinge/Maquet Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The CardioSave Hybrid/Rescue Intra-Aortic Balloon Pumps (IABP) are cardiac (heart) assist devices used with patients undergoing cardiac and non-cardiac surgery, and to treat adult patients who have acute coronary syndrome or complications of heart failure." | "Getinge/Datascope/Maquet is recalling these products due to complaints of fluid leaks. Fluid entering the Cardiosave IABP may cause unexpected pump shutdown or the inability to initiate therapy. There has been one death and 71 complaints reported about this device issue."
FDA MD Recalls | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device: 7Fr | "The Arrow-Trerotola Over-The-Wire (PTD) Kit Percutaneous Thrombolytic Device: 7FR includes a rotatable catheter with an outer sheath and an inner cable with a self-expanding basket. The Arrow-Trerotola Percutaneous Thrombolyic Device (PTD) Catheter is used with the Arrow Rotator Drive Unit, to remove clots in adult patients who have arterio-venous (AV) fistulas and synthetic dialysis grafts." | "Arrow International Inc, a subsidiary of Teleflex, is recalling the Arrow-Trerotola Over-The-Wire 7FR PTD Kits due to the risk of the orange inner lumen of the catheter's tip component separating from the basket. If the orange inner lumen separates from the basket, it may fracture and detach and block the blood vessel(s). If the orange inner lumen detaches from the basket, health consequences depend upon where the fractured tip component embolizes. If the embolization is local to the treatment target site, retrieval may be attempted, requiring an additional intervention and consequent delay of therapy. In some cases, the embolization could be central or possibly even to the heart or pulmonary arteries. This may lead to serious adverse events such as vessel damage, need for additional medical procedures, or possibly death. There have been seven complaints and no injuries or deaths reported for this device."
FDA MD Recalls | Ellume COVID-19 Home Test | clear from the name | "Ellume is recalling certain lots of the COVID-19 Home Test because they have higher-than-acceptable false positive test results for SARS-CoV-2. The reliability of negative test results is not affected. For these tests, a false positive test result shows that a person has the virus when they do not have it and could lead to: Delayed diagnosis or treatment for the actual cause of the person’s illness, which could be another life-threatening disease that is not COVID-19. Further spread of the SARS-CoV-2 virus when presumed positive people are grouped into cohorts (that is, they are housed together) based on false test results. The person receiving unnecessary COVID-19 treatment from a health care provider, such as antiviral treatment, convalescent plasma, or monoclonal antibody treatment, which can result in side effects. Disregard for the recommended precautions against COVID-19, including vaccination. Isolation, including monitoring household or close contacts for symptoms, limiting contact with family or friends, and missing school or work. There have been 35 reports of false positive results sent to the FDA and no deaths reported."
FDA MD Recalls | Custom Convenience Kits containing the Cardinal Health Monoject Flush Prefilled Syringe including: Leaderflex Insertion Kit with Ultrasound, Insertion Tray-RX, Universal Procedure Pack with Split Drape, Port Insertion Pack, and Procedure Pack | "Aligned Medical Solutions has packaged the Cardinal Health Monoject Flush Prefilled Syringe (0.9% Sodium Chloride) in a convenient manner for use in a general clinical procedure. All other components in the kit are not affected by this recall." | "Aligned Medical Solutions is recalling its custom convenience kits due to the potential for the plunger of the Cardinal Health Monoject Flush Prefilled Syringe (0.9% Sodium Chloride) to draw back after the air has been expelled and reintroduce air back into the syringe. If a clinician is not aware of air being reintroduced into the syringe, the clinician could inadvertently push air into the vascular system, creating the potential for an air embolism. The use of affected product may cause serious adverse health consequences, including death. There have been 42 reports of plunger retraction and introduction of air into the Cardinal Health Monoject Flush Prefilled Syringe (0.9% Sodium Chloride). There have been no reports of complaints, injuries or deaths related to this device issue."
FDA MD Recalls | Cardiosave Hybrid/Rescue Intra-Aortic Balloon Pump Battery Packs | "The CardioSave Hybrid/Rescue Intra-Aortic Balloon Pumps (IABP) are cardiac (heart) assist devices used with patients undergoing cardiac and non-cardiac surgery, and to treat adult patients who have acute coronary syndrome or complications of heart failure." | "Datascope/Getinge/Maquet is recalling this product due to the risk of the battery failing and having a shortened run-time due to substandard batteries not meeting performance specifications being released to customers, which may cause the device to stop working when operated by battery only. If a patient requires life-supporting therapy with an IABP and the device does not work or if therapy is stopped during use due to battery failure, the patient will be at risk of serious injury, including death. Both Cardiosave Hybrid and Rescue IABP monitors display battery life to the user, prompting intervention with low battery alarms when alternative power sources are indicated. Datascope/Getting/Maquet reports receiving six complaints, and no reports of injuries or deaths related to this issue. However, there is a potential for underreporting since the end user reporting a failed battery or short battery-run time cannot be aware that they originally received a substandard battery. "
FDA MD Recalls | ROSA One 3.1 Brain Application | "The ROSA One 3.1 Brain Application is a robotic platform that helps neurosurgeons with positioning medical instruments or implants during surgery. The device is composed of a compact robotic arm and a touch screen mounted on a stand. Different types of instruments or tools (for example, biopsy needles, stimulation or recording electrodes, endoscopes) may be attached to the end of the robotic arm, depending on the surgical procedure. " | "Zimmer Biomet is recalling this product due to a software error that could lead to incorrect placement of instruments during stereotactic neurosurgical procedures (for example, techniques to direct the tip of a tool using coordinates provided by medical imaging to reach a specific part of the brain). If this occurs, it could cause adverse events such as stroke, serious injury, severe disability, and death. There have been three complaints about this device issue. There have been no deaths or injuries reported about this device issue."
FDA MD Recalls | Software for Alinity m SARS-CoV-2 AMP Kit and Alinity m Resp-4-Plex AMP Kit | "The Alinity m SARS-CoV-2 Assay and the Alinity m Resp-4-Plex are laboratory tests used to detect RNA of SARS-CoV-2—the virus that causes COVID-19. The test uses swab samples taken from the back of the nose and throat (nasopharyngeal) or the back of the mouth and throat (oropharyngeal) or by a procedure that washes the lungs with saline to collect a sample (bronchoalveolar lavage) of people suspected to have an illness caused by the COVID-19 virus. Samples for both tests are taken at a health care setting and require special laboratory processing equipment to get results. The special processing equipment uses software to automate the mixing of the chemicals for the tests." | "Abbott Molecular, Inc. is recalling the software for the Alinity m SARS-CoV-2 AMP Kits and Alinity m Resp-4-Plex AMP Kits because of the potential to issue false positive results when being used to detect SARS-CoV-2. For these tests, a “false positive” is a test result that indicates that a person has the virus when they do not actually have it. The false positive results may be related to the current mixing parameters for the chemicals that are combined with patient samples to detect the viruses. In some cases, the mixing parameters may cause a sample in one well of the machine’s tray to overflow into another sample well. Overflow into the wells of true negative samples that are positioned near positive samples in the tray could produce false positive results. A false positive result for SARS-CoV-2 could lead to: Delayed diagnosis or treatment for the actual cause of the person’s illness, which could be another life-threatening disease that is not COVID-19. Further spread of the SARS-CoV-2 virus when presumed positive people are grouped into cohorts (that is, they are housed together) based on false test results. The person receiving unnecessary COVID-19 treatment from a health care provider, such as antiviral treatment, convalescent plasma, or monoclonal antibody treatment, which can result in side effects. People not taking recommended precautionary measures against COVID-19, including vaccination. Isolation measures, including monitoring of household or other close contacts for symptoms, limiting contact with family or friends, and missing school or work. There have been no deaths or adverse health consequences reported from use of this test."
FDA MD Recalls | DeRoyal Procedure Packs containing Smiths Medical NORMOFLO Irrigation Warming Set | "DeRoyal Industries’ procedure packs include an irrigation warming set that may be used with a fluid warming device. Fluid warming devices may be used by health care providers to: Provide more blood for patients suffering from blood loss due to trauma or surgery. Rewarm patients after surgery to prevent hypothermia. Irrigate during urology procedures." | "DeRoyal Industries, Inc. is recalling their procedure packs which contain the recalled Smiths Medical NORMOFLO Irrigation Warming Set because of the potential for aluminum ions to leach into warmed fluids. If this occurs, contaminated fluids can circulate through a patient’s body and the patient can be exposed to high levels of aluminum. Exposure to harmful levels of aluminum may not be easily recognized and the effects could cause serious adverse events including cognitive decline, dementia, nervous system damage, and even death. Patients who have symptoms may experience bone or muscle pain or weakness, anemia, seizures, or coma. There have been no complaints, injuries, or reports of death for this issue."
FDA MD Recalls | ZOOM 71 Reperfusion Catheter | "The ZOOM 71 Reperfusion Catheter is used with the ZOOM Aspiration Tubing and ZOOM Aspiration Pump (or equivalent vacuum pump) to remove blood clots in the brain of a patient within eight hours of an acute ischemic stroke caused by blocked or narrowed arteries." | "Imperative Care Inc. is recalling specific lots of this product due to increased risk of the Zoom 71 Reperfusion Catheter breaking at a certain point (distal tip) during use, such as when the catheter is retracted forcefully. If the device breaks during use, this could lead to fractured pieces left inside the patient’s bloodstream resulting in serious adverse events such as blockage of blood vessels, stroke, and death. There have been nine reports of serious injuries. There have been no deaths reported."
FDA MD Recalls | All MiniMed Remote Controllers (model MMT-500 and MMT-503) used with a Medtronic MiniMed 508 insulin pump or the MiniMed Paradigm family of insulin pumps | "People who have diabetes may use the MiniMed insulin pump to deliver insulin for the management of their diabetes. The pump system includes an optional remote controller device which is designed to communicate wirelessly with the pump to deliver a specific amount of insulin to the person with diabetes. The following table shows the recalled Medtronic remote controllers, used with the MiniMed 508 insulin pump or the MiniMed Paradigm family of insulin pumps. The remote controllers impacted by this issue are older models that use previous-generation technology. As of July 2018, Medtronic is no longer manufacturing or distributing these remote controllers." | "Medtronic is recalling all remote controllers used with either the MiniMed 508 insulin pump or the MiniMed Paradigm family of insulin pumps due to potential cybersecurity risks. An unauthorized person (someone other than a patient, patient caregiver, or health care provider) could potentially record and replay the wireless communication between the remote and the MiniMed insulin pump. Using specialized equipment, an unauthorized person could instruct the pump to either over-deliver insulin to a patient, leading to low blood sugar (hypoglycemia), or stop insulin delivery, leading to high blood sugar and diabetic ketoacidosis, even death. If you have never programmed a remote controller ID into the pump and never programmed the easy bolus option, you are not impacted by this vulnerability. To date, the FDA is not aware of any reports of patient harm related to these potential cybersecurity risks."
FDA MD Recalls | AMSORB PLUS PREFILLED G-CAN 1.0L | "AMSORB PLUS is a chemical substance used in anesthesia machines to absorb carbon dioxide (CO2) the patient breathes out while under anesthesia. This substance is provided in a prefilled canister (Prefilled G-CAN 1.0L) that is attached to the ventilator/anesthesia machine. " | "Armstrong Medical Limited is recalling this product due to reports that some AMSORB PLUS PREFILLED G-CAN 1.0L canisters have difficulties with gas flow that may cause a reduced flow of air to the patient. The difficulties in flow may cause a failure in the pre-use or “CHECK-OUT” test on the anesthesia machine. It may also occur when a patient is already under anesthesia if an affected canister is attached to the anesthesia machine while a patient is already under anesthesia. If the flow issue is not addressed, this may prevent the ventilator from providing enough breathing assistance to the patient (hypoventilation), leading to a build-up of carbon dioxide in the patient’s body and causing negative health effects. There have been no deaths or injuries reported for this recall."
FDA MD Recalls | Cordis SUPER TORQUE MB Angiographic Catheter with Radiopaque Marker Bands | "The Cordis SUPER TORQUE MB Angiographic Catheter with Radiopaque Marker Bands is a catheter used to visualize and measure parts of the vascular system when used with radiopaque (X-ray detectable) contrast media." | "Cordis Corporation is recalling the Cordis SUPER TORQUE MB Angiographic Catheter with Radiopaque Marker Bands due to the potential for the marker bands to move or dislodge during procedures. This can happen during procedures where the catheter is trapped between another device and the vessel wall. If this happens this could cause serious adverse events including delays in the procedure, the need for added medical procedures, heart attack, or stroke. There have been 167 complaints, 8 injuries, and no deaths reported for this issue. "
FDA MD Recalls | Pipeline Flex Embolization Device and Pipeline Flex Embolization Device with Shield Technology | "The Pipeline Flex Embolization Device and Pipeline Flex Embolization Device with Shield Technology are permanent mesh cylinders (stent) braided from platinum and tungsten and cobalt-chromium-nickel alloy wires. These devices are intended for the treatment of brain aneurysms that bulge or balloon out the sides of the blood vessel (wide-neck and fusiform). The Pipeline Flex Devices include a guidewire-based delivery system used to place the implant inside the patient." | "Medtronic (Micro Therapeutics, Inc. d/b/a ev3 Neurovascular) is recalling the Pipeline Flex Embolization Device and Pipeline Flex Embolization Device with Shield Technology because there is a risk of the delivery system’s wire and tubes fracturing and breaking off when the system is being used to place, retrieve, or move the stent inside a patient. In the United States (U.S.), the recall is only limited to the Pipeline Flex Embolization Device. Outside the U.S., the recall applies to both the Pipeline Flex Embolization Device and the Pipeline Flex Embolization Device with Shield Technology. Fractured pieces could be left inside the patient’s brain bloodstream. It is also possible that attempts to retrieve the fractured pieces may make the patient’s condition worse. The fragments can also cause other serious adverse health consequences such as continued blockage of blood vessels, stroke, and death. There have been 59 reported device malfunctions, 10 serious injuries, and two deaths related to this recall."
FDA MD Recalls | NORMOFLO Irrigation Fluid Warmer and NORMOFLO Irrigation Warming Set | "Fluid warming systems are used for warming blood, blood products, and intravenous solutions before use. There are several instances where fluid warmer products may be used, such as: When more blood is needed for patients suffering from blood loss due to trauma or surgery To rewarm patients after surgery to prevent hypothermia For irrigation in urology procedures Health care providers use fluid warming systems in a hospital or clinical setting. Fluid warmers may be used over a full range of flow rates from slow to rapid." | "Abbott Molecular, Inc. is recalling the software for the Alinity m SARS-CoV-2 AMP Kits and Alinity m Resp-4-Plex AMP Kits because of the potential to issue false positive results when being used to detect SARS-CoV-2. For these tests, a “false positive” i
FDA MD Recalls | ultrasound gel products and lotions manufactured by Eco-Med are marketed by a number of distributors, under different product names | "Eco-Med ultrasound gels and lotions are used during an ultrasound test to create images of organs and structures inside the body. Ultrasound gel is a clear, water-based gel that is applied to the skin on the area to be examined. The gel helps with the transmission of sound waves." | "All ultrasound gels and lotions manufactured by Eco-Med are being recalled due to risk of bacterial contamination with Burkholderia cepacia complex (Bcc). The FDA independently confirmed that distributed product tested positive for bacterial contamination. The use of affected ultrasound gels and lotions contaminated with Bcc may lead to serious infections, including bloodstream infections, which may result in sepsis or death. As of Aug. 31, 2021, there have been at least 66 infections, including 60 bloodstream infections associated with these affected products, per the Centers for Disease Control and Prevention (CDC)."
FDA MD Recalls | Alaris Infusion Pump Module Model 8100 Bezel purchased or installed by Bio-Medical Equipment Service Co. | "The Alaris Infusion Pump Module Model 8100 is used to deliver fluids such as medications, nutrients, blood, and other therapies into a person’s body, in controlled amounts. The bezel component is critical to the proper performance of the infusion pump. Bezels from Bio-Medical Equipment Service Co. are used to service and repair infusion pump modules." | "Bio-Medical Equipment Service Co. is recalling the affected products and devices because the bezel components may crack or separate, leading to inaccurate delivery of fluids to patients. The separation of one or more bezel posts may result in: Free flow of fluids to a patient Over-delivery or under- delivery of fluids delivered to a patient Interruption of fluids delivered to a patient If this device issue occurs, it could cause serious patient harm. There have been no deaths, complaints, or injuries reported for this recall. This recall is related to Tenacore’s recent recall of the Alaris pump bezel assembly and Alaris infusion pumps repaired with the bezel assembly."
FDA MD Recalls | Monoject Flush Prefilled Syringes (0.9% sodium chloride) | "Monoject Saline Flush Prefilled Syringes are used to flush out many different types of intravenous (IV) tubing and other devices that give health care workers access to the blood vessels (intravascular access devices) for fluid and medication delivery. These syringes are prefilled with preservative-free saline in various sizes and different amounts of saline (volumes). They are disposable and used one time only." | "Cardinal Health is recalling three models of Monoject Saline Flush Prefilled Syringes because it is possible the syringe’s plunger may pull air into the syringe after the health care provider has pushed out (expelled) the air. The materials used to make the syringe let the plunger slip back in the syringe. If air is injected into blood vessels, it can block the vessel (air embolism). Air embolisms can result in serious adverse health consequences, including stroke or death. There have been 37 complaints, no injuries, and no deaths reported for this issue."
FDA MD Recalls | Argyle UVC Insertion Tray containing Safety Scalpel N11 | "The Safety Scalpel N11 is a component included in the Argyle UVC Insertion Tray. Clinicians use the scalpel in the process of inserting an Umbilical Venous Catheter (UVC) into the umbilical artery or vein of newborns." | "Cardinal Health is recalling the Argyle UVC Insertion Tray because it does not include specific Instructions for Use (IFU) for the Safety Scalpel N11 contained in the tray. The Safety Scalpel N11 has a permanent locking feature that will not allow a clinician to unlock the scalpel once they have closed the safety shield and pushed it forward into a locked position (Figure 4). If the clinician is not aware of this permanent locking feature, there could be a delay in the procedure due to obtaining a replacement scalpel. Critically ill newborn patients may require emergency umbilical vein catheter placement. A clinician’s inability to use the scalpel (when in the permanent locked position) poses a safety risk to the patient that could result in delayed treatment which could lead to serious injury or death. There have been 10 complaints reported about this device issue. There were two reports of death in which the Safety Scalpel within the UVC kit locked permanently, and the clinician was unable to unlock for use during emergency UVC placement. No other injuries have been reported."
FDA MD Recalls | HeartWare HVAD System | The HeartWare Ventricular Assist Device (HVAD) System is used to help the heart continue to pump blood to the rest of the body. The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients for whom heart transplants are not planned. | "Medtronic is stopping the distribution and sale of the Heartware HVAD System because: There is an increased risk of neurological adverse events and mortality associated with the internal pump. If the internal pump stops, it may delay restarting or fail to restart. Both problems may lead to death or serious patient harm including stroke, heart attack, worsening heart failure, the need for additional procedures and hospitalizations. There is a growing body of observational clinical comparisons showing a higher frequency of neurological adverse events and mortality among HVAD System patients as compared to those who receive other commercially available durable left ventricular assist devices (LVAD). Medtronic reports there are over 100 complaints involving a delay or failure to restart of the HVAD internal pump, which led to a total of 14 deaths and 13 pump removals."
FDA MD Recalls | Dose IQ Safety Software (used with Spectrum IQ Infusion System) | "The Dose IQ Safety Software is a standalone computer-based software program that allows care providers to create a drug library of pre-programmed dosing information for specific medications. Those medications are delivered into a person’s body in controlled amounts using the Spectrum IQ infusion pump." | "Baxter Healthcare is recalling the Dose IQ Safety Software used with the Spectrum IQ Infusion System. A software defect may lead to mismatched drug information in drug libraries created using version 9.0.x of the computer-based software. If the software’s drug identifier does not match the drug information in the pump, the drugs may be improperly configured and cause serious adverse events including delay in therapy, under-infusion, or over-infusion of medication. There have been 15 complaints, no injuries, and no deaths reported for this issue. "
FDA MD Recalls | INGENIO Family of Pacemakers and CRT-Ps (includes models ADVANTIO DR EL, INGENIO DR EL and VITALIO DR EL) | "Boston Scientific INGENIO family pacemakers or cardiac resynchronization therapy pacemakers (CRT-Ps) are devices used with patients who have low heart rates, and patients with moderate to severe heart failure, a condition in which the heart cannot pump enough blood to meet the body's needs." | "Boston Scientific is recalling INGENIO family of pacemakers and CRT-Ps due to the risk of transitioning to safety mode. Safety mode is intended to provide backup if the device is faulty. However, in safety mode, there is a risk of inappropriate loss of pacing due to sensing of muscle contractions. If the device goes into safety mode, the device cannot be reprogrammed and must be replaced. The use of affected product may cause serious adverse health consequences, including early device replacement, loss of pacing or ability to regulate heart rate with serious or life threatening injury (for example, need of temporary pacing where a healthcare provider delivers controlled electric pulses to pace a heart), worsening of heart failure and death. There have been 65 reported incidents, including three injuries which required patients to receive temporary external pacing. There have been no reports of death."
FDA MD Recalls | Recirculator 8.0 Disposable Lavage Kit | "The Recirculator 8.0 Disposable Lavage Kit is used with the company’s Hyperthermic Perfusion System. The disposable lavage kit allows the perfusion system to circulate warm, sterile solution inside a patient’s chest (thoracic) or abdominal (peritoneal) cavity to reach a target temperature during surgery." | "Eight Medical is recalling the Recirculator 8.0 Disposable Lavage Kit because the disposable heating element within the kit may release aluminum into the fluid during warming therapy. When the fluid circulates through the patient’s body, the patient can be exposed to high levels of aluminum, which can be toxic and cause serious adverse events including death. There have been no complaints, injuries, or deaths reported for this issue."
FDA MD Recalls | V60 Plus ventilators and all V60 ventilators upgraded to enable High Flow Therapy (software version 3.00 and 3.10) | "V60 and V60 Plus ventilators equipped with high flow therapy are used to provide patients with breathing assistance at high concentration of oxygen at a higher flow than typical oxygen therapy. These devices include a design safety mechanism limiting the amount of flow that can be delivered to a patient in situations where the system pressure reaches a default maximum pressure limit, due to a partial obstruction in the breathing circuit. " | "Philips Respironics is recalling the V60 and V60 Plus ventilators that provide High Flow Therapy (Software Version 3.00 and Software Version 3.10) due to the risk to the patients who rely on the ventilator to provide high flow oxygen therapy if the system pressure reaches the default maximum limit. This can occur if the oxygen flow is partially blocked for any reason. If the issue causing maximum system pressure is not resolved by the health care provider, the ventilator will continue to provide the patient with lower oxygen flow rate and issue a low priority alarm. This could lead to serious adverse events, including death. There have been 61 incidents, 25 injuries, and no deaths reported for this issue."
FDA MD Recalls | Continuous Ventilator, Minimum Ventilatory Support, Facility Use: Philips Respironics E30 with Humidifier Continuous Ventilator, Non-life Supporting devices: DreamStation ASV, DreamStation ST, AVAPS, SystemOne ASV4, C-Series ASV, C-Series S/T and AVAPS, OmniLab Advanced+, and Noncontinuous Ventilators: SystemOne (Q-Series), DreamStation, DreamStation Go, Dorma 400, Dorma 500, and REMstar SE Auto) | "The Continuous and Non-Continuous Ventilators are used to provide invasive and non-invasive support for people needing respiratory support or treatment for sleep disorders. The type of ventilators used depends on the therapeutic needs of a patient. Depending on the device, the CPAPs, BiPAPs, and ventilators may be used in the home, hospital and other institutional settings. For more information, please see the UPDATE: Certain Philips Respironics Ventilators, BiPAP, and CPAP Machines Recalled Due to Potential Health Risks: FDA Safety Communication." | "Philips Respironics, Inc., is recalling its Continuous and Non-Continuous Ventilators due to the polyester-based polyurethane (PE-PUR) sound abatement foam, which is used to reduce sound and vibration in these affected devices, which may break down and potentially enter the device’s air pathway. If this occurs, black debris from the foam or certain chemicals released into the device’s air pathway may be inhaled or swallowed by the person using the device. The exposure to debris or chemicals could cause serious adverse events in patients such as irritation (skin, eye, and respiratory tract), inflammation, headache, asthma, hypersensitivity, nausea/vomiting, adverse effects to other organs (e.g., kidneys and liver) and toxic carcinogenic effects. There have been more than 1200 complaints and more than 100 injuries reported for this issue."
FDA MD Recalls | Trilogy 100, Trilogy 200, Garbin Plus, Aeris, LifeVent 200, BiPAP V30, BiPAP A30/Hybrid/A40 | "These devices are used to provide breathing assistance. Specifically: A continuous ventilator device is intended to mechanically control or assist patient breathing by delivering a predetermined percentage of oxygen in the breathing gas. A bilevel positive airway pressure (also known as BiPAP, BiLevel PAP, or BPAP) machine pumps air under pressure into the airway of the lungs. BiPAP machines have a higher pressure when you breathe in and lower pressure when you breathe out." | "Philips Respironics, Inc., recalled certain ventilators and BiPAP devices due to two issues related to the polyester-based polyurethane (PE-PUR) sound abatement foam that is used in Philips Continuous and Non-Continuous Ventilators, i.e., PE-PUR foam may break down into small pieces that may enter the device’s air pathway and then be swallowed or inhaled when the device is used. Additionally, the PE-PUR foam may release (off-gas) certain chemicals. The off-gassed chemicals and foam particles may lead to serious or life-threatening injuries, difficulty breathing (respiratory distress), swelling (inflammation), a lack of oxygen (hypoxia), too much carbon dioxide (hypercarbia), or toxic reactions. There have been 83 complaints, no injuries, and no deaths reported for these issues."
FDA MD Recalls | GENOSYL DS; Nitric Oxide Delivery System | "The GENOSYL DS; Nitric Oxide Delivery System is used to deliver a constant flow of GENOSYL (Nitric Oxide), which is a drug to help open blood vessels and improve oxygen levels in critically ill newborns with respiratory failure. The system consists of a cassette, which contains the drug, and a console. The device is used with a ventilator in hospital or healthcare settings." | "Vero Biotech is recalling its GENOSYL DS; Nitric Oxide Delivery System due to a software issue that leads to errors in the delivery of nitric oxide. Typically, this issue caused delivery of lower-than-expected dosage of nitric oxide during the transition between primary and backup console. If this happens, this could cause serious adverse events such as drops in oxygen level, heart problems, and clinical instability in the newborn. There have been 11 complaints, three injuries and no deaths reported for this issue."
FDA MD Recalls | Lyra SARS-CoV-2 Assay (M120) | "The Lyra SARS-CoV-2 Assay (M120) is a real-time polymerase chain reaction (RT-PCR) assay used to qualitatively detect nucleic acid from SARS-CoV-2, the virus that causes COVID-19. The test uses nasal, nasopharyngeal (NP), or oropharyngeal (OP) swab samples from patients suspected by their healthcare provider of illness caused by the COVID-19 virus. The SARS-CoV-2 virus is generally detectable in upper respiratory specimens during the onset of an infection. Positive results indicate the presence of SARS-CoV-2 RNA; however, patient history and other diagnostic information needs to be considered to determine if the patient has an infection. Positive results do not rule out bacterial infection or co-infection with other viruses." | "Quidel is recalling the Lyra SARS-CoV-2 Assay (M120) due to a significant risk of false negative results for patients with relatively high amounts of SARS‐CoV‐2 virus potentially causing the PCR amplification to occur before a cycle‐threshold (Ct) value ≤5 when using the following thermocyclers: ThermoFisher QuantStudio 7 Pro, Applied Biosystems 7500 Fast Dx, Applied Biosystems 7500, Bio-Rad CFX96 Touch, Roche LightCycler 480, or Qiagen RotorGene MDx False negative results may lead to delayed diagnosis or inappropriate treatment of SARS-CoV-2 that may cause patient harm, serious illness, and death. False negative results can also lead to further spread of the SARS-CoV-2 virus, when presumed negative patients are introduced into groups within health care, long-term care, or other similar facilities. Actions to limit exposure based on false negative results might not be taken, such as isolating infected individuals. Quidel has received five complaints about the Lyra SARS-CoV-2 Assay (M120), however, there have been no reports of injuries or death from this issue."
FDA MD Recalls | Angiographic Guidewire Component | "Medtronic Vascular’s Angiographic Guidewire Component is a guidewire used during angiography or other interventional procedures to help place catheters into the vasculature. " | "Medtronic Vascular is recalling the Angiographic Guidewire Component because devices were not sterilized before being shipped directly to hospitals. If patients are exposed to the non-sterile device, serious adverse events could occur such as infection, sepsis, and death. There have been two complaints, and no reported injuries or deaths related to this issue. However, there is potential for underreporting as physicians may not have been aware that devices were non-sterile."
FDA MD Recalls | LeadCare II, LeadCare Plus, and LeadCare Ultra Blood Lead Tests | "The LeadCare II, LeadCare Plus, and LeadCare Ultra Blood Lead Tests are used to find out a person's blood lead level. The test uses a finger or heel stick whole blood (capillary) sample. After the sample is mixed with treatment reagent, a LeadCare analyzer measures the amount of lead collected on a sensor. Magellan's LeadCare II is a point-of-care (CLIA-waived) blood lead testing system. The LeadCare systems are used in clinical laboratories, doctor's offices, clinics, and hospitals throughout the United States." | "Magellan Diagnostics, Inc., is recalling its LeadCare II, LeadCare Plus, and LeadCare Ultra Blood Lead Tests due to a significant risk of falsely low results. The FDA has significant concerns that the performance of the test may provide falsely low results and may lead to health risks in special populations such as young children and pregnant individuals. A pregnant or lactating individual's exposure to lead is concerning because it not only may cause health problems for the parent but can result in lead exposure to the developing baby. Obtaining falsely low results may lead to patient harm including delayed puberty, reduced postnatal growth, decreased IQ, and inattention and behavior problems in children."
FDA MD Recalls | Jelco Hypodermic Needle-Pro Fixed Needle Insulin Syringes | "The Jelco Hypodermic Needle-Pro Fixed Needle Insulin Syringe is a one-time use syringe used to draw and inject a common formulation of insulin (U-100) into a patient." | "Smiths Medical is recalling the affected products because these insulin syringes may have graduated markings that are printed incorrectly on the syringe barrel. Specifically, the “odd numbered” graduation markings on the syringe barrel may be skewed approximately 20 degrees upward. If the skewed markings are used to measure an insulin dose, patients may receive too much or too little insulin. This overdose or underdose of insulin can lead to serious patient harm, including death. There have been nine complaints, no injuries, and no deaths reported for this issue."
FDA MD Recalls | Adult Chest Tube Tray Fistula On-Off Kit Fistula (On/Off All in One bag) Dialysis Access | "Avid Medical’s medical convenience kits include a set of devices that are used to complete routine medical and surgical care such as: Draining fluid from an adult’s chest Surgical repair of an abnormal connection between an artery and a vein (arteriovenous fistula) Dressing changes around the surgical site of an arteriovenous fistula Creating a dialysis access where a patient is connected to the dialysis machine Each convenience kit listed above includes a Chloraprep™ 3 mL applicator, which is used to sanitize skin before surgery or catheter procedures. The Chloraprep 3 mL applicator is manufactured by another firm, BD/Carefusion 213." | "Avid Medical is recalling medical convenience kits that include the BD/Carefusion Chloraprep™ 3mL applicator. The applicator was recalled due to the risk of contamination with a specific type of fungus called Aspergillus penicillioides. If skin preparation products are contaminated with Aspergillus penicillioides, the fungus can cause serious systemic infection, sepsis, illness, and death to the patient. If the fungus is introduced in the patient’s bloodstream during placement of an intravascular catheter, the catheter may need to be removed, requiring additional medical procedures. If the fungus infects a surgical site, the patient may require medical and surgical treatments and require long-term treatment with antifungal drugs. There have been no deaths, complaints, or reported injuries related to this issue."
FDA MD Recalls | DeRoyal Pacemaker Tray Pgybk DeRoyal Heart Cath Procedure Pack DeRoyal Angio Cath Removal Tray | "DeRoyal Industries’ surgical procedure packs include a set of devices that are used for routine cardiac procedures such as heart catheterizations, pacemaker-related procedures and x-ray heart imaging (angiogram) procedures. Each pack listed above includes Hospira Pfizer 1% lidocaine, a local anesthetic that prevents or reduces pain by stopping nerves from sending pain signals during a surgical or catheter procedure." | "DeRoyal is recalling the surgical procedure packs because the packs contain 1% lidocaine that has been mislabeled as 0.5% bupivacaine. Though both lidocaine and bupivacaine are local anesthetics, their dosing is different. If 1% lidocaine is given to the patient instead of 0.5% bupivacaine, the patient may be underdosed and experience pain during the procedure. If 0.5% bupivacaine is given to the patient instead of 1% lidocaine, it may cause an overdose of bupivacaine with potential life threatening or fatal consequences. There have been no deaths, complaints, or reported injuries related to this issue. This recall is related to a Hospira recall for the same mislabeling issue."
FDA MD Recalls | Alaris Infusion Pump Module Model 8100 Bezel, purchased and/or installed by Infusion Pump Repair | "The Alaris Infusion Pump Module Model 8100 is used to deliver fluids such as medications, nutrients, blood, and other therapies into a person’s body, in controlled amounts. The bezel component is critical to the proper performance of the infusion pump. Bezels from Infusion Pump Repair are used to service and repair infusion pump modules." | "Infusion Pump Repair is recalling the affected products and devices because the bezel components may crack or separate, leading to inaccurate delivery of fluids to patients. The separation of one or more bezel posts may result in: free flow of fluids to a patient over delivery or under delivery of fluids delivered to a patient interruption of fluids delivered to a patient If this device issue occurs, this could cause serious patient harm. There have been no deaths, complaints or injuries reported for this recall. This recall is related to Tenacore’s recent recallExternal Link Disclaimer of the Alaris pump bezel assembly and Alaris infusion pumps repaired with the bezel assembly."
FDA MD Recalls | Innova SARS-CoV-2 Antigen Rapid Qualitative Test (also distributed under the names Innova COVID-19 Self-Test Kit (3T Configuration), Innova SARS-CoV-2-Antigen Rapid Qualitative Test (7T Configuration), and Innova SARS-CoV-2-Antigen Rapid Qualitative Test (25T Configuration)) | "The Innova SARS-CoV-2 Antigen Rapid Qualitative Test claimed to determine if a person has an active COVID-19 infection. The test used a nasal swab sample and test strip to detect specific proteins, called antigens, from the SARS-CoV-2 virus. If the nasal sample had SARS-CoV-2 antigens, a colored test line should have appeared on the test strip, indicating that the person may have COVID-19. If the nasal sample did not have SARS-CoV-2 antigens, a colored line should not have appeared on the test strip. The test has not been authorized, cleared, or approved by FDA for commercial distribution in the United States." | "Innova Medical Group is recalling its SARS-CoV-2 Antigen Rapid Qualitative Test. Labeling distributed with certain configurations of the test includes performance claims that did not accurately reflect the performance estimates observed during the clinical studies of the tests. The performance characteristics of the test have not been adequately established, presenting a risk of false results. False-negative results may lead to delayed diagnosis or inappropriate treatment of SARS-CoV-2, which may cause patient harm including serious illness and death. False-negative results can also lead to further spread of the SARS-CoV-2 virus, including when presumed negative patients are grouped into cohorts in health care, long-term care, and other facilities based on false test results. False-positive results could lead to a delay in the correct diagnosis and the initiation of an appropriate treatment for the actual cause of patient illness, which could be another life-threatening disease that is not SARS-CoV-2. False-positive results could also lead to further spread of the SARS-CoV-2 virus when presumed positive patients are grouped into cohorts based on false test results."
FDA MD Recalls | Alaris Infusion Pump Module Model 8100 Bezel, purchased and/or installed by The Biomed Guys | "The Alaris Infusion Pump Module Model 8100 is used to deliver fluids such as medications, nutrients, blood, and other therapies into a person’s body, in controlled amounts. The bezel component is critical to the proper performance of the infusion pump. Bezels from The Biomed Guys are used to service and repair infusion pump modules." | "The Biomed Guys is recalling the affected products and devices because the bezel components may crack or separate, leading to inaccurate delivery of fluids to patients. The separation of one or more bezel posts may result in: Free flow of fluids to a patient Over delivery or under delivery of fluids delivered to a patient Interruption of fluids delivered to a patient If this device issue occurs, this could cause serious patient harm. There have been no deaths, complaints or injuries reported for this recall. This recall is related to Tenacore’s recent recallExternal Link Disclaimer of the Alaris pump bezel assembly and Alaris infusion pumps repaired with the bezel assembly."
FDA MD Recalls | Alaris Infusion Pump Module Model 8100 Bezel, purchased and/or installed by Step-Har Medical | "The Alaris Infusion Pump Module Model 8100 is used to deliver fluids such as medications, nutrients, blood, and other therapies into a person’s body, in controlled amounts. The bezel component is critical to the proper performance of the infusion pump. Bezels from Step-Har Medical are used to service and repair infusion pump modules." | "Step-Har Medical is recalling the affected products and devices because the bezel components may crack or separate, leading to inaccurate delivery of fluids to patients. The separation of one or more bezel posts may result in: free flow of fluids to a patient over delivery or under delivery of fluids delivered to a patient interruption of fluids delivered to a patient If this device issue occurs, this could cause serious patient harm. There have been no deaths, complaints or injuries reported for this recall. This recall is related to Tenacore’s recent recallExternal Link Disclaimer of the Alaris pump bezel assembly and Alaris infusion pumps repaired with the bezel assembly."
FDA MD Recalls | Lepu Medical Technology SARS-CoV-2 Antigen Rapid Test Kit and Leccurate SARS-CoV-2 Antibody Rapid Test Kit (Colloidal Gold Immunochromatography) | The Lepu Medical Technology SARS-CoV-2 Antigen Rapid Test Kit uses a nasal swab sample to detect proteins, called antigens, from the SARS-CoV-2 virus. The SARS-CoV-2 Antigen Rapid test is used to find out if a person has an active COVID-19 infection. "The Leccurate SARS-CoV-2 Antibody Rapid Test Kit (Colloidal Gold Immunochromatography) looks for antibodies in a person’s immune system produced in response to SARS-CoV-2, the virus that causes COVID-19, and should not be used to diagnose an active COVID-19 infection. The Leccurate SARS-CoV-2 Antibody Rapid Test Kit (Colloidal Gold Immunochromatoraphy) uses serum, plasma, or blood samples." | "Lepu Medical Technology is recalling the SARS-CoV-2 Antigen Rapid Test Kit and Leccurate SARS-CoV-2 Antibody Rapid Test Kit (Colloidal Gold Immunochromatography) due to likely risk of false results. Neither test was authorized, cleared, or approved for marketing or distribution in the United States by the FDA. There have been no reports of injuries or death from use of these tests."
FDA MD Recalls | Medical Action Industries, Inc. 306 Medical Convenience Kits | "Medical Action Industries’ Medical Convenience Kits include a set of devices that are used to complete routine medical care such as: Dressing changes around a central line, a catheter in a large vein that delivers fluids and allows blood draws Injections Midline catheter insertion to deliver medications or fluids into a vein Dressing changes for left ventricular assist devices, a type of heart pump Collection of blood cultures to test for infections Each convenience kit listed above includes a Chloraprep™ 3 mL applicator, which is used to sanitize skin prior to wound care, catheter procedure or blood collection. The Chloraprep 3 mL applicator is manufactured by another firm, BD/Carefusion 213." | "Medical Action Industries’ is recalling the Medical Convenience Kits that include the BD/Carefusion Chloraprep™ 3mL applicator. This component was recalled due to the risk of contamination with a specific type of fungus called Aspergillus penicillioides. If skin preparation products are contaminated with Aspergillus penicillioides, the fungus can cause serious systemic infection, sepsis, illness, and death to the patient. If the fungus is introduced in the patient’s bloodstream during placement of an intravascular catheter, the catheter may need to be removed, requiring additional medical procedures. If the fungus infects a surgical site, the patient may require medical and surgical treatments and require long-term treatment with antifungal drugs. There have been no deaths, complaints, or reported injuries related to this issue."
FDA MD Recalls | VICI VENOUS STENT System (VICI SDS) and VICI RDS VENOUS STENT System | "The VICI SDS and VICI RDS VENOUS STENT Systems are intended for the treatment of obstructions and occlusions in the narrowed or blocked venous veins." | "Aligned Medical Solutions is recalling its custom convenience kits due to the potential for the plunger of the Cardinal Health Monoject Flush Prefilled Syringe (0.9% Sodium Chloride) to draw back after the air has been expelled and reintroduce air back i
FDA MD Recalls | Scalable Bradycardia Platform (SBP) Pacemakers: Assurity™ and Endurity™ | "Assurity™ and Endurity™ are implantable pacemakers that detect when the heart is beating too slowly (bradycardia) and then send signals to the heart to make it beat at the correct pace. These pacemakers can be used to provide pacing for one chamber of the heart or both chambers, based on the patient’s condition." | "Abbott (formally known as “St. Jude Medical”) is recalling a subset of Assurity and Endurity pacemakers built using specific manufacturing equipment, that were then distributed from April 2015 to February 2019. A small number of devices from that time frame have experienced problems when moisture is able to get inside the device. The moisture can cause an electrical short, that may lead to: A loss of device pacing Telemetry failure or errors in information Early and fast battery drain Less time between the first battery depletion warning (elective replacement indicator or ERI) and the device’s end of service (EOS) If the device is unable to deliver pacing, patients may experience slow or irregular heartbeat, fainting, shortness of breath, tiredness, dizziness, or discomfort. Additionally, shorter battery life and device life may lead to an additional pacemaker replacement procedure sooner than expected. Finally, if the system does not relay accurate information via telemetry, medical providers may not know to provide treatment. There have been 135 complaints, 135 injuries, and no deaths reported for this issue."
FDA MD Recalls | HeartWare HVAD System Instructions for Use and Patient Manual that accompany components including HeartWare HVAD Pump, Controller, Driveline Cable, Driveline Cover, and Surgical Tool Kits | "The HeartWare Ventricular Assist Device (HVAD) System helps the heart continue to pump blood to the rest of the body. The HVAD System is used as a bridge to heart transplants in patients who are at risk of death from end-stage left ventricular heart failure, to give the heart tissue time to recover, or as a final, or destination, therapy for patients where new transplants are not planned." | "Medtronic is recalling their HeartWare HVAD System to provide updated Instructions for Use (IFU) and Patient Manual (PM) due to safety issues with (1) Carrying Cases, (2) Driveline Cover Orientation and; (3) Controller Power-Up Sequence. Carrying Case: A drop of the carrying case, caused by damage to the case or improper wear, can disconnect the driveline and then interrupt pump power. The IFU and PM will be updated to provide information about the lifespan for the carrying cases, and information about how to safely clean and wear the Convertible Patient Pack. Driveline Cover Orientation: When the driveline cover is first installed during the surgical implant procedure, if installed properly it will always cover the driveline connector. However, if a patient removes the cover during a controller exchange, there is a possibility that the driveline cover could be put on backwards. In that reversed orientation, the fit of the cover inadvertently causes the driveline locking mechanism to be in the unlocked position, which could cause temporary or accidental driveline disconnects. The PM will be updated to inform patients to keep the driveline cover on when disconnecting and reconnecting the driveline during a controller exchange. Controller Power-Up Sequence: During power-up, the LED lights turn red and can be misunderstood as a “red alarm,” leading to an unnecessary controller exchange. The IFU and PM will be updated to make clear that the expected power-up sequence causes the alarm indicator LEDs and both sets of battery LEDs to turn red for 2.5 seconds while the controller LCD displays the power-on message. If using the HVAD system and (1) the carrying case breaks and the driveline pulls out of the controller as it drops, or (2) the driveline disconnects from backwards driveline cover orientation; or (3) a controller exchange is performed unnecessarily due to confusion of start-up behavior as a “red alarm” battery failure, this may cause serious patient harm, including death. There has been 1 death and 64 injuries reported to the FDA for these issues."
FDA MD Recalls | Alaris Infusion Pump Module Model 8100 Front Bezel, purchased and/or installed by Pacific Medical Group (DBA Avante Health Solutions) | "The Alaris Infusion Pump Module Model 8100 is used to deliver fluids such as medications, nutrients, blood, and other therapies into a person’s body, in controlled amounts. The front bezel component is critical to the proper performance of the infusion pump. Bezel posts from Pacific Medical Group are used to service and repair infusion pump modules." | "Pacific Medical Group (DBA Avante Health Solutions) is recalling the affected products and devices because the front bezel components may crack or separate, leading to inaccurate delivery of fluids to patients. The separation of one or more bezel posts may result in: free flow of fluids to patient over delivery or under delivery of fluids delivered to a patient interruption of fluids delivered to a patient There have been 62 complaints reported about this device issue and one death may have been related to this recall. No other injuries have been reported. This recall is related to Tenacore’s recent recall of the Alaris pump bezel assembly and Alaris infusion pumps repaired with the bezel assembly. There have been 15 complaints, no injuries, and no deaths reported for this issue. "
FDA MD Recalls | Precise PRO Rx US Carotid System | "The Precise PRO Rx US Carotid System is used to treat patients with narrowed carotid arteries. The system includes a metal (nitinol) self-expanding stent preloaded on a delivery catheter used to place the stent." | "Cordis Corporation is recalling their Precise PRO Rx US Carotid System because of a risk of separation of the atraumatic distal tip of the sheathed delivery system in patients. If the device separates during use this may cause serious adverse events such as removal of the separated tip from the carotid artery, embolization distally, or stroke. There have been 7 complaints, including 5 reported injuries about this device issue. No deaths have been reported."
FDA MD Recalls | Medtronic Bio-Console 560 Extracorporeal Blood Pumping Console (Bio-Console 560) | "The Bio-Console 560 Extracorporeal Blood Pumping Console is used to pump and return blood to a person during cardiopulmonary bypass, a procedure that temporarily takes over the function of the heart and lungs during surgery. The cardiopulmonary bypass procedure provides both circulatory and breathing support to a person whose heart and lungs are too weak to pump enough blood to the rest of the body. The Bio-Console 560 can pump blood for up to six hours each time used." | "Medtronic is recalling the Bio-Console 560 Extracorporeal Blood Pumping Console due to reports of electrical failure causing the device to stop. Other issues observed by users include the interface going blank, smoke coming from the device and a burning odor during use. In cases where the pump stopped or the interface went blank, users had to operate the pump using a hand crank until the device restarted or could be replaced by a backup unit. If the pump stops, patients could experience: Insufficient blood flow potentially leading to organ dysfunction Blood clots potentially leading to low blood pressure Reversible or irreversible neurological dysfunction Death There have been nine complaints reported about this device issue, including three injuries. No deaths have been reported."
FDA MD Recalls | Alaris Infusion Pumps serviced by Tenacore | "The Alaris Pump is used to deliver fluids such as medications, nutrients, solutions, blood, and other therapies into a person’s body in controlled amounts. Bezel posts by Tenacore are used to service and repair infusion pump modules." | "Tenacore is recalling the Alaris Pump because the bezel repair posts may crack or separate leading to inaccurate delivery of fluids to patients. The separation of one or more bezel repair posts may result in: free flow of fluids to patient over delivery or under delivery of fluids delivered to a patient interruption of fluids delivered to a patient device leaking device component burning or smoking power supply failure or malfunction If this device issue occurs, this could cause serious patient harm. There have been two complaints reported about this device issue. No injuries or deaths have been reported."
FDA MD Recalls | BD Alaris Infusion Pump Module | "The BD Alaris Pump Module System is an infusion pump and vital signs monitoring system that consists of a PC Unit, the Guardrails Suite MX, and up to four removable infusion or monitoring modules (channels). The System is used to deliver fluids including medications, blood, or blood products, into a patient's body in a controlled manner. It is used in hospitals and other health care facilities." | "CareFusion 303 Inc is recalling the Alaris Pump Module because of the risk of the keypad lifting up due to fluid entry. This could lead to keys that become unresponsive or stuck (See Figure 1). This may lead to an infusion delay or interruption or prevent clinicians from changing fluid or medication infusions on the affected devices. High-risk patient populations who are receiving life-sustaining infusions are at the greatest risk of harm. For these patients, delays or interruption of infusions can cause serious injury or death. There have been 79 complaints regarding this device issue. There are no reports of injuries or death."
FDA MD Recalls | HeartWare HVAD Battery Cables, Data Cables, Adapter Cables and Controller 2.0 Ports | "The HeartWare Ventricular Assist Device (HVAD) Battery Cables, Data Cables, Adapter Cables and Controller 2.0 Ports are part of the HeartWare HVAD System, which is used to help the heart continue to pump blood to the rest of the body (See Figure 1). The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients where new transplants are not planned." | "Medtronic is recalling the HeartWare HVAD Battery Cables, Data Cables, Adapter Cables and Controller 2.0 Ports because of risk of wear and tear of the connector plugs (power sources, data cable, and alarm adapter) which could cause damage to the Controller port metal pins (for example, bent pins). Damaged controller ports may prevent power cables and data cables from being connected to the controller and lead to a full or partial stop of the pump. This could cause serious patient harm including loss of consciousness, hospitalization, heart attack, or death. There have been 12 deaths and 8 injuries reported to the FDA."
FDA MD Recalls | ThermaCor 1200 Disposable Sets | "The ThermaCor 1200 Rapid Thermal Infusion System disposable sets are part of the ThermaCor 1200 Rapid Thermal Infusion System, which is used for fluid or bolus delivery. The system is made of a footswitch for hands-free fluid control, a fluid delivery device, and a single-use disposable set with supply lines that can be used with intravenous (IV) bags or surgical equipment." | "Smisson-Cartledge Biomedical is recalling their ThermaCor 1200 Rapid Thermal Infusion System Disposable Sets because a part of the ThermaCor 1200 Disposable Set, the cassette (See Figure 1), which warms fluids directly with an aluminum plate may leak aluminum into the fluids and expose patients to high levels of the metal. The recall is specific to the disposable cassette portion of the device, not to the full pump. Use of the affected product could cause increase exposure to aluminum ions. Exposure to high levels of aluminum ions could cause serious patient harm such as bone or muscle pain and weakness, anemia, seizures, or coma. There have been no complaints, reports of injuries, or deaths related to this device issue."
FDA MD Recalls | Medtronic Evera, Viva, Brava, Claria, Amplia, Compia, and Visia Implantable Cardioverter Defibrillators (ICDs) and Cardiac Resynchronization Therapy (CRT-Ds) | "Medtronic’s Evera, Viva, Brava, Claria, Amplia, Compia, and Visia Implantable Cardioverter Defibrillators (ICDs) and Cardiac Resynchronization Therapy (CRT-Ds) are devices that provide pacing for slow heart rhythms, and electrical shock or pacing to stop dangerously fast heart rhythms. ICDs and CRT-Ds are both implanted under the skin in the upper chest area with connecting insulated wires called "leads" that go into the heart. A person may need an ICD or CRT-D if their heartbeat is too slow (bradycardia), too fast (tachycardia), or needs coordination to treat heart failure." | "Medtronic is recalling the specified ICDs and CRT-Ds due to an unexpected and rapid decrease in battery life. The decrease in battery life is caused by a short circuit and will cause some devices to produce a “Recommended Replacement Time” (first warning that the battery is low) earlier than expected. Some devices may progress from “Recommended Replacement Time” to full battery depletion within as little as one day. If the user does not respond to the first warning, the device may stop functioning. The likelihood that this issue will occur is constant after approximately three years after device use. There have been 444 complaints regarding these devices. There are 264 Medical Device Reports, with 18 injuries including people experiencing bradycardia (slow heart rhythm) or heart failure symptoms. Zero deaths have been reported."
FDA MD Recalls | Medtronic Valiant Navion Thoracic Stent Graft System | "The Valiant Navion Thoracic Stent Graft System is designed to repair lesions of the descending thoracic aorta, located in the body’s largest artery (aorta) which passes through the lower part of the chest. A surgeon uses a long tube-like device (catheter) to place the stent graft inside the aorta. Once placed, the stent expands to fit within the aorta to provide a new path for blood to flow from the heart to the lower part of the body." | "Medtronic is recalling the Valiant Navion Thoracic Stent Graft System due to stent fractures and endoleak concerns. Following device implantation, patients may experience stent ring enlargement beyond design specification, stent fractures, holes in the graft fabric (Type III endoleaks), life-threatening bleeding, aortic rupture, or death. There have been 11 reported issues related to this device including two injuries and one death. There have been four cases of Type III endoleak, four cases of stent facture, and seven cases of stent ring enlargement. Some patients experienced multiple findings."
FDA MD Recalls | ACIST Kodama Intravascular Ultrasound Catheter | "The ACIST Kodama Intravascular Ultrasound Catheter (HD-IVUS) uses high-frequency sound waves (ultrasound), to view the inside of an artery, including the structure of the artery, and any devices such as stents. It is inserted through an artery in the leg and guided to the area that health care providers need to see. This catheter is intended to be used with the ACIST HDi System." | "ACIST Medical Systems is recalling their ACIST Kodama Intravascular Ultrasound Catheter because the O-ring housing tubing may squeeze and damage the O-ring (see Fig. 1). If there is damage, pieces of the O-ring may break free and flush into the patient’s artery during use. This could cause serious patient harm including vessel blockage, heart attack, heart arrhythmia, or stroke, with associated chest pain or discomfort. Patients with heart ventricles that do not function properly are at higher risk of experiencing serious harm due to sudden blockage of blood flow. Additionally, if the HD-IVUS catheter is imaging an artery that supplies blood to another partly blocked or narrowed artery, there is a higher risk of blockages in more than one artery. There have been no complaints, injuries, or deaths reported."
FDA MD Recalls | Valkyrie LTOWB Collection and Administration Low Titer Type O FWB Transfusion Set and Low Titer Type O Donor Collection Set Fresh Whole Blood Transfusion Set and Fresh Whole Blood Donor Set | "The convenience kits under recall are used by military medical personnel to get and transfuse donor blood to patients on the battlefield. Each kit has a blood bag with a needle attached and includes other components to screen blood." | "Combat Medical is recalling the convenience kits because the needle in the blood pack may bend or disconnect from the blood bag. The possible needle damage only occurs during the packaging process before being received by the user. However, a user will not be able to identify if the needle is damaged until the pack has been opened for use. A bent/disconnected needle could potentially result in: Bruising or hematoma (an area of clotted blood under the skin). Delay, change, or inability to treat patients by giving needed blood, which may lead to injury or death. There have been four complaints about this device issue. There have been no reports of injuries or deaths."
FDA MD Recalls | HVAD Pump Implant Kits for the HeartWare HVAD System | "The HeartWare Ventricular Assist Device (HVAD) Pump Implant Kit is part of the HeartWare HVAD System, which is used to help the heart continue to pump blood to the rest of the body. The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients where new transplants are not planned." | "Medtronic is recalling the HVAD Pump Implant Kit because the device may fail to initially start, restart, or have a delay in restarting after the pump was stopped. These delays or failures to start or restart have occurred during preimplant testing, during the implant, or in a variety of post-implant situations. If the device has delays or fails to start or restart, this could cause serious patient harm including a heart attack, worsening heart failure, the need for additional procedures and hospitalizations, or death. There have been 29 complaints about this device issue, which include 19 serious injuries and 8 cases of patients who had a life-threatening event but recovered without long term effects. Two deaths have been reported."
FDA MD Recalls | Hillrom Liko Multirall 200 Overhead Lift | "The Liko Multirall 200 Overhead Lift is a general-purpose lift to move patients from room to room. This lift is part of the the Multirall 200 overhead lift system which has an overhead lift motor, the S65 rail carriage hook, and a Q-link strap. The multirall system is used in health care settings like nursing homes, rehabilitation facilities and hospitals." | "Hillrom is recalling the Liko Multirall 200 Overhead Lift due to customer reports that the Q-link strap lock does not attach to the S65 carriage hook as it should. If the strap lock does not attach, the motor or the patient may fall. Use of the affected product may cause adverse events such as serious injury and death. There have been 34 complaints about this device issue and 22 reports of serious injuries. Two deaths have been reported."
FDA MD Recalls | EMBLEM S-ICD (Subcutaneous Implantable Cardioverter Defibrillator) | "The EMBLEM S-ICD (Subcutaneous Implantable Cardioverter Defibrillator) is part of the Boston Scientific S-ICD System. The Boston Scientific S-ICD is an implantable cardioverter defibrillator that is intended to provide electrical shock to stop dangerously fast heart rhythms and pacing for a short time after shocks if needed. These cardiac devices are placed under the skin in the upper chest area." | "Boston Scientific is recalling the EMBLEM S-ICD because a manufacturing process may allow moisture to get inside the defibrillator and cause a short-circuit when it tries to deliver high voltage shocks. If this happens during use, patients may experience less shock than intended or may not receive a shock at all. The device may also beep, not respond to a device check in, and issue battery alerts. A defibrillator with the manufacturing error may delay or prevent the device from delivering a lifesaving electrical shock to a person in cardiac arrest (tachycardia) and lead to serious injury or death. Additional surgeries may also be needed to replace failed devices. The manufacturer has received six complaints about this device issue. There have been no reports of injuries or deaths."
FDA MD Recalls | EMBLEM S-ICD Subcutaneous Electrode | "The EMBLEM S-ICD Subcutaneous Electrode Model 3501 is part of the Boston Scientific S-ICD System. The Boston Scientific S-ICD is an implantable cardioverter defibrillator that is intended to provide electrical shock to stop dangerously fast heart rhythms and pacing for a short time after shocks if needed. These cardiac devices are implanted under the skin in the upper chest area." | "Boston Scientific is recalling the EMBLEM S-ICD Subcutaneous Electrode because of increased risk of fractures at a specific point (distal to the proximal sense ring) shown in Figure 1. If the device fractures during use, it could become unable to deliver therapy to slow very fast heartbeats from cardiac arrest (tachycardia). A failed device may cause serious adverse events. Examples include injury or death if cardiac arrest cannot be treated or need for additional surgery to replace failed devices. There have been 27 complaints about this device issue and 26 reports of serious injuries. One death has been reported."
FDA MD Recalls | Penumbra JET 7 Reperfusion Catheter with Xtra Flex Technology, also known as JET 7 Xtra Flex | "As a component of the Penumbra System with the Penumbra Aspiration Pump and Penumbra Aspiration Tubing, the JET 7 Xtra Flex catheter and JET 7MAX configuration (JET 7 Xtra Flex catheter and MAX Delivery Device) are medical devices intended to restore blood flow by removing clots using continuous aspiration in patients experiencing an acute ischemic stroke within 8 hours of symptom onset who are ineligible for intravenous tissue plasminogen activator (IV t-PA) or who fail IV t-PA therapy. This recall does not apply to the Penumbra JET 7 Reperfusion Catheter with Standard Tip." | "On December 17, 2020, Penumbra recalled the Penumbra JET 7 Xtra Flex because the catheter may become susceptible to distal tip damage during use. The FDA has received over 200 medical device reports (MDRs) associated with the JET 7 Xtra Flex catheter, including deaths, serious injuries, and malfunctions. Twenty of these MDRs describe 14 unique patient deaths, which include reports from different reporting sources for a single adverse event. Other MDRs describe serious patient injury, such as vessel damage, hemorrhage, and cerebral infarction. Device failure modes reported in the MDRs include ballooning, expansion, rupture, breakage or complete separation, and exposure of internal support coils near the distal tip region of the JET 7 Xtra Flex catheter."

## FDA MDR 2022
Base | Device name | Device use | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
FDA MD Recalls | Mighty Bliss electric heating pads | "Mighty Bliss electric heating pads produce heat through electrical elements. These devices are flexible and may include heat-retaining material such as gel, fluid, or vegetable matter to help distribute the heat evenly across the heating pad surface. They are used for temporary relief of aches and pains." | "Whele LLC is recalling Mighty Bliss electric heating pads after receiving customer complaints related to the product burning customers' skin, overheating, sparking, burning, or presenting other electrical problems. Use of the recalled Mighty Bliss electric heating pads could lead to serious adverse events including infections, scar tissue from a skin contraction, skin irritation, skin burns, electric shock, and fire. Whele LLC reported 286 customer complaints between July 2021 and September 2022 related to this recall. There have been 31 reported injuries such as electric shocks, burns, and skin burns, rashes or irritation. There are no reports of deaths related to this issue."
FDA MD Recalls | Arrow AutoCAT 2, AutoCAT 2 Refurbished, AeroautoCAT 2, AutoCAT 2 Wave, AutoCAT 2 Wave refurbished, AeroautoCAT 2 Wave, AeroautoCAT 2 Wave refurbished, AC3 NA/EMEA, AC3 Optimus NA/EMEA, AC3 Optimus NA/ EMEA, AC3 Optimus Refurbished, AC3 Optimus NA/ AJLA | "The Arrow Intra-Aortic Balloon Pumps (IABP) are cardiac (heart) assist devices used with patients undergoing cardiac and non-cardiac surgery, and to treat adult patients who have acute coronary syndrome or complications of heart failure." | "Arrow International, LLC, a subsidiary of Teleflex, Inc, is recalling Arrow AutoCAT 2 and AC3 Intra-Aortic Balloon Pumps after receiving complaints that battery power is not lasting as long as expected. A fully charged battery should last 90 minutes, however, multiple users report the devices are experiencing shorter run times. The battery life warnings, designed to sound at 20, 10, and 5 minutes of battery power remaining, are also inaccurate as the battery depletes more rapidly than indicated. If the pump’s battery depletes, the pump may stop when operated on battery power only. Sudden stopping of the pump’s mechanical support to the patient’s heart can lead to cardiac arrest, stroke, or death. There have been 241 complaints, with 135 reported pump stops. No deaths have been reported related to this issue. "
FDA MD Recalls | Arrow MAC Two-Lumen Central Venous Access Kits; Arrow Pressure Injectable Arrowg+ard Blue Plus Three-Lumen Central Venous Catheter (CVC) Kits | "The Arrow MAC Two-Lumen Central Venous Access Kit is intended to permit short term (less than 30 days) venous access and catheter introduction to the central circulation. The Arrow Pressure Injectable Arrowg+ard Blue Plus Three-Lumen CVC Kit is intended to provide short-term (less than 30 days) protection against catheter-related bloodstream infections. It is not intended to be used as a treatment for existing infections." | "Teleflex and their subsidiary Arrow International, LLC are recalling the Arrow MAC Two-Lumen Central Venous Access Kits; Arrow Pressure Injectable Arrowg+ard Blue Plus Three-Lumen Central Venous Catheter (CVC) Kits for the risk of a cross-lumen leak caused by inadequate connections between the top and bottom housings of the Micro Clave Clear Connectors included in the kits. Use of affected devices may cause bleeding, fluid leakage, delayed treatment, infection, air in the blood vessels (air embolism), other serious injuries or death. Teleflex/Arrow International LLC has reported no injuries or deaths related to this issue. "
FDA MD Recalls | Thermo Scientific Gram Negative IVD AST Sensititre Plate | "The Thermo Scientific Gram Negative IVD AST Sensititre Plate is an in vitro diagnostic product for clinical susceptibility testing of non-fastidious Gram-negative organisms, comprising of Enterobacterales, Pseudomonas aeruginosa, and other non-Enterobacteriaceae. Each plate contains various dilutions of antimicrobial agents at appropriate concentrations. Results can be read manually by visual reading of growth or automatically on an ARIS/OptiRead using fluorescence. The Thermo Scientific Gram Negative IVD AST Sensititre Plates are used only by trained laboratory personnel to test the susceptibility of Gram-negative bacterial infections to specific antibiotics by monitoring the detection of bacterial growth. Results are used as an aid to select the drug of choice for treatment. " | "Remel, Inc. is recalling the Thermo Scientific Gram Negative IVD AST Sensititre Plates due to the risk of potential false susceptible results for Proteae Tribe (including Proteus spp., Providencia spp., Morganella Morganii) when tested with the following antibiotics: Carbapenem (doripenem, ertapenem, meropenem, imipenem), cefepime, piperacillin/tazobactam, and aztreonam. The test results from the affected product may indicate that treatment with these antibiotics will stop or slow the growth of Proteae Tribe bacteria (including Proteus spp., Providencia spp., Morganella Morganii) when that may not be accurate. A false susceptibility result could lead patients to receive an ineffective treatment and experience serious injury or health consequences, including increased time in the hospital, unnecessary tests or procedures, treatment failure, sepsis, and even death. The firm has received one complaint associated with these products and FDA has received five Medical Device Reports associated with these products. There are no known injuries or deaths."
FDA MD Recalls | Dewei Medical Equipment Co. DNA/RNA Preservation Kits | "Dewei Medical Equipment Co. DNA/RNA Preservation Kits are inactivated viral transport media, used for the collection, preservation, and transportation of clinical samples that will be tested for viruses that may include SARS-CoV-2 (the virus that causes COVID-19), influenza, avian influenza, hand, foot, and mouth disease, measles, norovirus, and rotavirus. " | "Dewei Medical Equipment Co. is recalling DNA/RNA Preservation Kits because these products were distributed to U.S. customers without authorization, clearance, or approval from the FDA. This means there is a risk for false negative, false positive, or misinterpretation of results if these products are used to detect SARS-CoV-2 or other viruses. Use of these products may cause serious adverse health consequences or death. There have been no complaints, reports of injuries, or deaths associated with the use of this product. "
FDA MD Recalls | WatchCare Incontinence Management System (IMS) | "The WatchCare Incontinence Management System (IMS) includes specific Centrella, Progressa, and VersaCare hospital beds and disposable incontinence pads. It is designed to discreetly alert the caregiver of an incontinence event. The system is usually used in critical care and medical/surgical settings as well as in other clinical areas. " | "Baxter Hillrom is recalling the WatchCare IMS after receiving reports that the radiofrequency emissions (RF) from WatchCare devices may interfere with other medical devices, including equipment that is critical to the health and wellbeing of patients, such as: Infusion pumps Insulin pumps Blood glucose sensors Fetal monitors / dopplers Telemetry devices Bladder scanners Other third-party medical devices may also be affected. The RF interference from WatchCare may cause erroneous readings or malfunctions of these other devices on both patients and staff, which could lead to inappropriate medical treatments or lack of treatment. If WatchCare RF emissions affect a medically necessary device such as a blood glucose sensor, insulin pump, fetal monitor, or general infusion pump, among others, this issue could lead to serious injury or death. Baxter reported 96 complaints of interference. To date, interferences in all cases but two occurred at distances less than 1 meter. There is insufficient data about distance on the remaining two reported interferences. There have been no reports of serious injuries or deaths related to this issue. "
FDA MD Recalls | Insulet Omnipod DASH Insulin Management System Personal Diabetes Manager (PDM) | "The Omnipod DASH Insulin Management System is used to deliver insulin at set and variable rates to manage diabetes in people who need insulin. It includes two main parts: The wearable and tubeless pump which delivers the insulin, called the Pod, and a remote controller called the Personal Diabetes Manager (PDM). The PDM is a small, reusable handheld rechargeable battery-powered locked-down Android mobile device with features similar to a smartphone. The Pod is controlled exclusively using the PDM and the two parts interact wirelessly using secure Bluetooth." | "Insulet is recalling the Omnipod DASH Insulin Management System’s Personal Diabetes Manager (PDM) after receiving reports of PDM battery issues, including: Battery swelling Fluid leakage from the battery Extreme overheating that may pose a fire hazard Users could be exposed to battery fluid and extreme heat, including the potential for an explosion and/or fire, which could lead to serious injury or death. Insulet reports 455 complaints involving these battery issues, including three fires. There have been no reports of injuries or deaths. "
FDA MD Recalls | Iso-Gard Filter S | "Iso-Gard Filter S filters are breathing circuit bacterial filters that are connected to respiratory equipment or breathing systems in intensive care units and operating rooms. These filters protect the equipment and the patients from potential airborne contaminants. " | "Teleflex is recalling several types of the ISO-Gard Filter S after receiving reports that the filters may split or detach from breathing systems. If the filter splits or detaches during use, the filter may not work properly or efficiently. The issue may cause a patient to receive an insufficient supply of air and for the patient’s oxygen levels to drop below safe levels (desaturation), as well as potential cross-contamination. Use of these devices may cause serious injury or death. Teleflex reported that there have been 36 complaints and 4 injuries, including desaturation, related to this issue. There have been no reports of deaths or long-term injuries. "
FDA MD Recalls | Amara View Full Face Mask DreamWisp Nasal Mask DreamWear Full Face Mask Wisp and Wisp Youth Nasal Mask Therapy Mask 3100 NC/SP | "The Amara Full Face, DreamWisp Nasal, DreamWear Full Face, Wisp and Wisp Youth Nasal, and Therapy Mask 3100 masks are worn by a person using bilevel positive airway pressure (also known as Bilevel PAP, BiPAP, or BPAP) machines and continuous positive airway pressure (CPAP) machines to support breathing. The recalled masks are for single patient use in the home or multi-patient use in the hospital or other clinical environment. The recalled Philips masks may be used with other manufacturers’ BiPAP and CPAP machines. Users of any BiPAP or CPAP machine should check to see whether their mask is one of the recalled Philips masks. " | "Philips Respironics (Philips) is recalling certain Amara Full Face, DreamWisp Nasal, DreamWear Full Face, Wisp and Wisp Youth Nasal, and Therapy Mask 3100 masks due to a serious safety concern. The recalled masks have magnetic headgear clips to hold them in place. The magnets can potentially cause injury or death if people who use them, or people near a person using a recalled mask, have certain implanted metallic medical devices or metallic objects in the body. The FDA issued a safety communication, Certain Philips Respironics Masks for BiPAP, CPAP Machines Recalled Due to Safety Issue with Magnets That May Affect Certain Medical Devices about this topic which includes more information on the types of medical devices and objects that may be affected. As of September 9, 2022, there have been 43 complaints, 14 of those were serious injuries, associated with this issue. The FDA is not aware of any associated deaths. "
FDA MD Recalls | COVID-19 Ag Rapid Test Device | "The COVID-19 Ag Rapid Test Device uses patient samples collected by nasal, nasopharyngeal, or oropharyngeal swab to detect proteins, called antigens, found on the SARS-CoV-2 virus. These tests are intended to be used by health care professionals." | "Jiangsu Well Biotech Co., Ltd. is recalling COVID-19 Ag Rapid Test Devices because they were distributed to U.S. customers without authorization, clearance, or approval from the FDA. The company has not provided the FDA with adequate validation data to show that the test's performance is accurate. This means there is a risk of potential false negative, false positive, or misinterpretation of results. Use of these tests may cause serious adverse health consequences or death. There is also a risk of injury if users follow any labeling instructions directing self-collection of nasopharyngeal or oropharyngeal samples. There have been no complaints or reports of injuries or deaths."
FDA MD Recalls | LifeSPARC System | "The LifeSPARC System is intended to pump the blood through an extracorporeal circuit for periods lasting less than 6 hours for the purpose of providing either: Full or partial cardiopulmonary bypass (i.e., circuit includes an oxygenator) during open surgical procedures on the heart or great vessels; or Temporary circulatory bypass for diversion of flow around a planned disruption of the circulatory pathway necessary for open surgical procedures on the aorta or vena cava. The LifeSPARC System has two components: the LifeSPARC Pump and the LifeSPARC Controller. The LifeSPARC Pump is a single-use pump, and the LifeSPARC Controller provides the interface between pump and user, as well as the power and electrical signals to drive the pump. " | "LivaNova (TandemLife) is recalling the LifeSPARC Controller, part of the LifeSPARC System due to a software malfunction. The controller’s monitoring feature may erroneously detect frozen or unresponsive software and trigger the device to enter Critical Failure mode—clearing the controller screen and issuing an alarm that cannot be muted or turned off. While the LifeSPARC Pump should continue to run at the set speed with the allowance of manual speed adjustment in Critical Failure mode, the user has to replace the controller using instructions from the Operations Manual, before disconnecting the pump from the frozen controller. If the user does not follow these specific instructions and powers off the frozen controller prior to acquiring and setting up the backup controller, the pump may stop for an extended period of time during the replacement process. An extended pump stop during use may cause serious injury or death to the patient. LivaNova reported that there have been 66 complaints, and the FDA has received two injuries and no deaths due to this issue."
FDA MD Recalls | Philips Respironics BiPAP A30, A40, V30, and OmniLab Advanced+ | "A bi-level positive airway pressure (also known as Bilevel PAP, BiPAP, or BPAP) machine pumps air under varying pressure into the airway of the lungs. BiPAP (Ventilator) machines use a higher pressure when you breathe in and lower pressure when you breathe out. These devices are prescribed for people with obstructive sleep apnea syndrome to keep their airways open during sleep, for those who have respiratory insufficiency (not enough oxygen breathed in or carbon dioxide breathed out), or respiratory failure (stop breathing). BiPAP machines may be used to treat adult and pediatric patients at home and in clinical environments, such as hospitals and sleep laboratories, depending on the instructions for use for the particular device model. " | "Philips Respironics (Philips) is recalling certain BiPAP machines that may contain a plastic contaminated with a non-compatible material. If that plastic is in the device motor, it may release certain chemicals of concern called volatile organic compounds (VOCs). The plastic may also cause the machine to fail and stop working suddenly during use. The potential risks associated with this recall include: Headache Dizziness Irritation in the eyes, nose, respiratory tract (airway), and skin Hypersensitivity reaction, such as an allergic reaction or another immune system reaction Nausea or vomiting If the plastic causes the machine to fail and stop working suddenly, it may also lead to serious injury or death. The FDA is not aware of any reports of serious injury or death associated with the recalled BiPAP machines due to the contaminated plastic issue."
FDA MD Recalls | Clearlink Basic Solution Set with Duovent | "The Clearlink Basic Solution Set with Duovent is used as part of a system to administer drugs and solutions to patients. The majority of the Clearlink Basic Solution Sets with Duovent are used for the delivery of hazardous drugs (chemotherapy)." | "Baxter Healthcare Corporation is recalling Clearlink Basic Solution Set with Duovent after increased customer reports of leaks. As the majority of the Clearlink Basic Solution Sets with Duovent are used for the delivery of hazardous drugs (chemotherapy), leakage could expose healthcare personnel, patients, and others to potentially hazardous drugs that may be toxic and/or are irritants. These leaks may also allow air into the set or breach the sterile fluid pathway, thereby increasing the risk of air embolism and contaminated infusions, respectively. Patients may suffer delayed or interrupted therapy or may not receive the necessary amount of their medication. These issues could lead to serious injury or death. There have been 83 complaints with no injuries and no deaths associated with this issue. "
FDA MD Recalls | NIM CONTACT Reinforced EMG Endotracheal Tube and the NIM Standard Reinforced EMG Endotracheal Tube | "The NIM CONTACT Reinforced EMG Endotracheal Tube and the NIM Standard Reinforced EMG Endotracheal Tubes are used during surgery to: Provide an airway for patient ventilation and Monitor electromyogram (EMG) activity and the nerve integrity of the thyroarytenoid muscle of the larynx." | "Medtronic is recalling the NIM CONTACT Reinforced EMG Endotracheal Tube and the NIM Standard Reinforced EMG Endotracheal Tubes after receiving customer complaints about obstruction of the endotracheal tube while in use in patients. If the tube obstructs, ventilation failure can occur. If the tube does not ventilate properly or obstructs the airway, patients may suffer oxygen deprivation, brain damage, or death. Medtronic reported that they have received 15 complaints, three injuries, and two deaths associated with this issue between March 31, 2020, and March 31, 2022. "
FDA MD Recalls | CereLink ICP Monitor | "The CereLink ICP Monitor is an intracranial pressure monitor. It monitors pressure in brain tissue, which is important to track in patients with head injuries, stroke, surgical and post-operative neurosurgical patients, and patients with other conditions. The device’s sensor is implanted in the brain and then connected by a wire to an external patient monitor that displays the pressure readings. These readings are used to monitor and when indicated to guide treatment. " | "Integra is recalling the CereLink ICP Monitor after reports that the device may display incorrect values for a patient’s intracranial pressure and display out-of-range pressure readings. If the device is not working properly, the patient may have to undergo additional brain surgeries to replace it, which involves the risks of infection, bleeding, and damage to tissue. Additionally, because the device is used to inform decision-making in critically ill patients (including people with severe traumatic brain injury and stroke) and managing intracranial pressure is key to their care, a malfunctioning device creates a risk of serious injury or death. Integra reported that there have been 105 global complaints associated with this recall as of July 31, 2022. As of August 24, 2022, the FDA is aware of at least 68 Medical Device Reports (MDRs) associated with this device problem, including reports of patient injury and one patient death. The patient death reported in the MDR described a malfunctioning CereLink ICP Monitor during use in a critically injured patient, which was mitigated by replacing the ICP sensor. The cause of patient death was determined by Integra to be unrelated to the CereLink ICP Monitor malfunction."
FDA MD Recalls | HAMILTON-C6 Intensive Care Ventilator | "The HAMILTON-C6 Intensive Care Ventilator is intended to provide mechanical ventilation or breathing support for infants (neonatal), children (pediatric), and adult patients. It is able to provide ventilator support invasively through a tube in the mouth, nose, or airway (trachea) or non-invasively through a mask placed over the mouth, nose, or face. The HAMILTON-C6 ventilator is used by qualified, trained healthcare personnel under the direction of a physician in hospitals and healthcare settings." | "Hamilton Medical AG is recalling the HAMILTON-C6 Intensive Care Ventilator after customer complaints revealed a hardware issue with the ventilator’s status indicator board. The status indicator board may become loose, allowing liquid to enter (ingress) between the indicator board and the ventilator’s main board. This water ingress may cause the ventilator to have a technical fault and revert to a safety ventilation mode or revert to an ambient state, which means the patient breathes ambient room air with no assistance or support from the machine. When the ventilator enters ambient state, an alternative source of ventilation must be provided immediately or the patient may experience interrupted ventilation, brain injury caused by lack of oxygen to the brain (hypoxia), buildup of carbon dioxide in the blood (hypercarbia), other serious injuries, or death. There have been 128 complaints (worldwide) and no injuries or deaths associated with this issue."
FDA MD Recalls | Intera 3000 Hepatic Artery Infusion Pump | "The Intera 3000 Hepatic Artery Infusion Pump is used to continuously deliver medicine through the hepatic artery, which supplies blood and oxygen to the liver and several other organs. The pump is implanted into adult patients and delivers chemotherapy medications to treat liver tumors. People with the Intera 3000 Hepatic Artery Infusion Pump are able to get back to daily activities and return to the treating physician for pump refills every two weeks." | "Intera Oncology is recalling certain units of Intera 3000 Hepatic Artery Infusion Pumps (50 products in the U.S.) after receiving reports from clinicians that the affected pumps were delivering medications (flow rate) faster than expected. If the pump delivers infusions at faster than expected flow rates, the patient may receive more medication than planned, which could result in increased risk of liver toxicity if the fast flow is not recognized and adjusted to deliver the appropriate dose. If too much drug is administered, it may result in serious adverse health consequences. There have been three incidents reported related to this issue, with no injuries or deaths."
FDA MD Recalls | Medtronic HVAD Batteries | "The HeartWare Ventricular Assist Device (HVAD) System is used to help the heart continue to pump blood to the rest of the body. The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients for whom heart transplants are not planned. The system operates using power from either AC or DC electricity or batteries." | "Medtronic, Inc. is recalling HeartWare HVAD System batteries because the batteries may experience electrical faults that cause them to unexpectedly fail. When this occurs batteries with an electrical fault may be unable to power the controller, unable to accept charge from the battery charger, or appear to remain charged when in use. If the battery fails and the patient is unable to replace the failed battery with a charged, working battery or with AC or DC power, the HVAD may stop working, leading to serious injury or death. Medtronic reports 1,159 complaints, six injuries, and one death related to this issue."
FDA MD Recalls | Cobalt XT, Cobalt, and Crome Implantable Cardioverter Defibrillators (ICDs) and Cardiac Resynchronization Therapy Defibrillators (CRT-Ds) | "Medtronic Cobalt/Crome Implantable Cardioverter Defibrillators (ICDs) and Cardiac Resynchronization Therapy Defibrillators (CRT-Ds) are implantable cardiac devices that are intended to monitor and regulate heart rate and rhythm. These devices can deliver a pacing impulse to correct a slow heartbeat (bradycardia), provide biventricular pacing/cardiac resynchronization therapyExternal Link Disclaimer, or produce an electric shock (cardioversion or defibrillation) to restore a regular heartbeat if a life-threatening arrhythmia occurs." | "Medtronic is recalling Cobalt/Crome Implantable Cardioverter Defibrillators (ICDs) and Resynchronization Therapy Defibrillators (CRT-Ds) after receiving reports of devices with a short circuit protection (SCP) alert resulting in reduced-energy electric shock delivery, instead of delivering a second phase of high voltage therapy. A reduced-energy electrical shock may fail to correct an irregular heartbeat (arrythmia) or may cause an irregular heartbeat. The harms associated with a reduced-energy electric shock or an inaccurate response to an SCP alert may cause serious injury or death. There have been 27 complaints, no injuries, and no deaths associated with this issue."
FDA MD Recalls | BD Intraosseous Needle Set Kits, BD Intraosseous Manual Driver Kits, BD Intraosseous Powered Drivers | "The BD Intraosseous Infusion System is used by clinicians to obtain access to blood vessels through a patient’s bone (intraosseous) when emergency, urgent, or medically necessary care is needed but it is difficult or impossible for the clinicians to obtain typical access through the blood vessels (intravenously). This access can then be used to inject or infuse medications/fluids. It is used in both adults and children for up to 24 hours." | "Becton Dickinson is recalling the BD Intraosseous Needle Set Kits, BD Intraosseous Manual Driver Kits, and BD Intraosseous Powered Drivers for three separate issues: The stylet may be difficult to separate from the needle or may not separate at all. This may result in inadvertent removal of the entire needle assembly during placement, or the clinician may be unable to remove the stylet from the needle, both resulting in functional loss of intraosseous access. The needle safety mechanism may not properly deploy as the stylet is removed from the intraosseous needle after placement. Metal discs may stick in the powered driver, which may render it unusable. These issues can cause delays in care due to the inability to place functional intraosseous access. Since intraosseous access is most often used in critically ill patients, including those with cardiopulmonary arrest or severe shock, the potential delays in care can cause serious injury or death. There is also the risk of needle stick injury. There have been 37 complaints with no serious injuries or deaths reported associated with these recall issues."
FDA MD Recalls | Viral Transport Media Containers | "Haimen Shengbang Laboratory Equipment Co. Ltd.’s Viral Transport Media Container is a sampling swab and a tube containing viral sampling medium (the liquid that maintains a specimen sample while it is transported to a laboratory for testing). This container and swab are used to collect specimens for transport to a laboratory for viral testing to detect infection, including testing for viruses such as SARS-CoV-2, the virus that causes COVID-19." | "Haimen Shengbang Laboratory Equipment is recalling Viral Transport Media Containers because these products were distributed to U.S. customers without authorization, clearance, or approval from the FDA. This means there is a risk for false negative, false positive, or misinterpretation of results if these products are used to detect SARS-CoV-2. Use of these products may cause serious adverse health consequences or death. There have been no complaints, reports of injuries, or deaths associated with the use of this product."
FDA MD Recalls | Oral Rapid SARS-CoV-2 Rapid Antigen Test Kit | "The Oral Rapid SARS-CoV-2 Rapid Antigen Test Kit uses patient samples collected by nasal swab to detect proteins, called antigens, found on the SARS-CoV-2 virus. These tests are intended to be used by clinical laboratories or health care workers for point-of-care testing." | "North American Diagnostics is recalling these tests because they were distributed to U.S. customers without authorization, clearance, or approval from the FDA. North American Diagnostics did not provide the FDA with adequate validation data to show that the test's performance is accurate. This means there is a risk of potential false negative, false positive, or misinterpretation of results. Use of these tests may cause serious adverse health consequences or death. There have been no complaints or reports of injuries or deaths."
FDA MD Recalls | Palindrome and Mahurkar Hemodialysis Catheters | "The Palindrome and Mahurkar catheters are implanted and used in hemodialysis (treatment to filter waste and water from your blood), apheresis (a technology that separates donated blood components to treat certain illnesses), and infusion (a method of putting fluids into the bloodstream)." | "Covidien, LLC (Medtronic) is recalling the Palindrome and Mahurkar catheters due to a catheter hub defect that will connect both extension catheters. There is a potential leaking condition within the hub of specific chronic dialysis catheters. This may be noticed when flushing one extension tube and the flow of fluid through the tip of the catheter returns unanticipated fluid through the adjacent extension tube. During treatment, this leak could result in mixing of the arterial and venous blood and lead to increased recirculation and poor dialysis, and the development of thrombi and emboli. The use of the defective catheter may cause serious adverse health outcomes, including bleeding, surgical removal and replacement of the affected catheter. There has been one complaint. There have been no reports of death or injury. "
FDA MD Recalls | Abacus Order Entry and Calculation Software | "Abacus is a software application that performs calculations for compounding liquid doses of medications. When connected to a compounder that mixes these doses according to the calculation, the Abacus software can translate a physician’s order into a compounded solution ready to be given to a patient. " | "Baxter Healthcare Corporation is recalling the Abacus software application due to a risk that final printed bag labels for compounded mixtures may contain incorrect information. The issue may occur if a user unintentionally or incorrectly modifies a label template that is used to provide clinical care. Incorrect values or patient names on final printed bag labels for compounded medications may cause serious harm, especially if the wrong amounts or wrong medications are administered to high-risk patients. There have been 5 complaints, no injuries, and no deaths associated with this software issue. "
FDA MD Recalls | Medfusion Syringe Infusion Pumps | "Smiths Medical Medfusion 4000 and 3500 Syringe Infusion Pumps are used to give fluids to patients in precisely controlled amounts. They deliver blood or blood products, lipids, drugs, antibiotics, enteral feedings and other therapeutic fluids through infusion tubing into a patient’s vein or through other cleared routes of administration. Syringe pumps are primarily used in the neonatal and pediatric populations or in operating rooms and intensive care units for the adult population." | "Smiths Medical is recalling Medfusion 3500 and 4000 Syringe Infusion Pumps for eight software malfunctions that affect different serial numbers and software versions. These malfunctions may cause serious harm or death to patients from under- or over-infusion, or delays in the delivery of critical medications to patients. The eight software issues are as follows: (1) False alarm for Primary Audible Alarm (PAA) system failure (2) Unanticipated Depleted Battery Alarms (3) Abnormal circuit board behavior, which may cause internal clock system failure (4) Intermittent Volume Over Time (IVOT) delivery mode where the infusion continues after system failure (5) Unanticipated clearance of Program Volume Delivered (PVD) (6) False alarm for Rate Below Recommended Minimum for Syringe Size (7) Incorrect bolus or loading dose time display (8) Network configuration may affect pump communications. Smiths Medical states there have been a total of 7 serious injuries and one death reported related to these issues. The Customer NotificationExternal Link Disclaimer identifies the injuries and/or deaths associated with each software issue. "
FDA MD Recalls | COVID Test Kit Nonsterile and Clean Catch Urine Kit | "The American Contract Systems (ACS) COVID Test Kit Nonsterile (a nasopharyngeal sample collection tray for COVID-19 testing, labeled as “COVID TEST KIT NONSTERILE 184312, AHCV18A”) includes a nasopharyngeal swab and Universal Transport Media and is intended to collect samples that are used to detect infection with SARS-CoV-2, the virus that causes COVID-19. The ACS Clean Catch Urine Kit is used to collect urine for further testing, including to test for potential urinary tract infection (UTI)." | "American Contract Systems is recalling the COVID Test Kit Nonsterile and Clean Catch Urine Kit products because they were assembled in an uncontrolled facility by people without proper training. As a result, the company is unable to verify that the kits will perform as expected. These kits have the potential to give false negative or false positive results or lead to misinterpretation of test results. Use of these affected products could cause serious adverse health consequences and death. The company has not received any complaints or reports of injuries or deaths associated with the use of these kits. "
FDA MD Recalls | Flow-c and Flow-e Anesthesia Systems | "The Getinge Flow-c and Flow-e Anesthesia Systems are intended to provide inhalation anesthesia and control the ventilation of patients who have no ability to breathe. These systems can also support the breathing of people with limited ability to do so themselves. It is used for patients ranging in age from newborns (neonatal) to adults and operated by trained anesthesiologists in hospital settings." | "Getinge is recalling the Flow-c and Flow-e Anesthesia Systems after receiving reports of cracked or broken on/off switches on the system’s suction unit. If the on/off switch breaks, the suction unit will not work to remove fluids such as phlegm, blood, or stomach (gastric) content from the patient’s mouth and/or airways. This issue could cause a delay during a procedure or in between procedures if the anesthesiologist needs to find another way to suction fluids. Additionally, if the suction device is not working for a longer period of time, these fluids could block a patient’s breathing. The risks of such a blockage may include: Choking Inability of oxygen to get into the blood (acute respiratory failure or ARDS) An infection in the lungs (pneumonia) that prevents oxygen from getting to the blood (respiratory failure) Brain injury caused by lack of oxygen to the brain (hypoxia) Death There have been 21 complaints, no injuries and no deaths associated with the use of this device."
FDA MD Recalls | CARESCAPE R860/Engstrom Carestation/Engstrom PRO Ventilator backup batteries and replacement backup batteries manufactured on or after April 1, 2019. | "The CARESCAPE R860 ventilator is intended to provide mechanical ventilation or breathing support to infants (neonatal), children (pediatric), and adult patients weighing 0.25 kg and above. The CARESCAPE R860 ventilator uses main AC power via a wall plug to operate. However, there is also a backup battery that is intended to keep the ventilator operating for a certain amount of time if a main AC power supply is not available, for example during a patient transport. Replacement backup batteries are available from the company to use when the original backup battery expires." | "GE Healthcare is recalling CARESCAPE R860 ventilator backup batteries, including replacement backup batteries, manufactured on or after April 1, 2019, because the batteries may run out before they are expected to do so. If the ventilator is running on battery power only when the battery fails, ventilation could stop completely, preventing the patient from receiving oxygen and breathing support. Lack of oxygen (hypoxia), especially if it occurs over a long period, can cause serious injury and death. There have been 1,553 complaints, no injuries, and no deaths associated with the use of this device."
FDA MD Recalls | Medtronic HVAD Batteries | "The HeartWare Ventricular Assist Device (HVAD) System is used to help the heart continue to pump blood to the rest of the body. The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients for whom heart transplants are not planned." | "Medtronic, Inc. is recalling a single lot of HeartWare HVAD System batteries due to welding defects that allow separation of the two cell battery packs used to power the system. The welding defect may cause the battery to malfunction and no longer provide power or prevent the battery from holding a full charge or properly recharging. As a result, the battery may fail suddenly. If this malfunction occurs, it will trigger a Power Disconnect alarm on the controller screen and in the Alarm Log tab of the HVAD Monitor while the affected battery is still connected to the controller. These alarms may temporarily resolve, but the battery will continue to fail. The use of these batteries may lead to device malfunction, death, or serious injury. Medtronic reports one death associated with this recall, and two complaints in the affected lot."
FDA MD Recalls | Volara system with in-line ventilator adaptor (OPTIMUS Handset 2) or Volara patient circuit kit (OPTIMUS OLE AC Patient Circuit Kit) | "The Volara system is intended to help people clear mucus out of their airways, expand the lungs, and to treat or prevent a partial collapsed lung (pulmonary atelectasis). The system’s in-line ventilator adaptor component makes it possible for it to be used together with a ventilator (in-line) in home-care settings. " | "Baxter Healthcare Corporation, and its subsidiary company Hillrom, are recalling the Volara system because the in-line ventilator adaptor may prevent home-use patients from getting enough oxygen from their ventilators. The risks to affected patients include: Choking on mucus or other airway secretions An infection in the lungs (pneumonia) that prevents oxygen from getting to the blood (respiratory failure) Brain injury caused by lack of oxygen to the brain (hypoxia) Death The risk of serious injury or death is more significant in home-care settings if the caregivers are not trained properly, the device is not connected properly, or if the caregiver is not prepared to address any issues that may arise caused by use of this device. There has been one complaint and one injury, as well as two deaths, associated with the use of this device."
FDA MD Recalls | SafeStar 55 Breathing System Filter | "The Draeger, Inc SafeStar 55 is a breathing system filter that is used as part of a ventilator when a patient is under anesthesia or needs breathing assistance. The SafeStar 55 Breathing System Filter is intended to reduce the contaminants like bacteria, other micro-organisms, and small particles, that reach a patient who is receiving breathing support from the ventilator." | "Draeger, Inc is recalling a specific lot (LT2103) of the SafeStar 55 Breathing System Filter because a manual inspection process led to some defective filters, including some that may be partially obstructed, to be inadvertently distributed instead of destroyed. If the filter on a ventilator or breathing system is obstructed, oxygen may not flow properly to the patient. That lack of oxygen (hypoxia) can have serious effects including death. There has been one complaint and one injury associated with the use of this device. There have been no reported deaths."
FDA MD Recalls | Oral Rapid SARS-CoV-2 Antigen Rapid Test Kit and Joysbio SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold) | "The Oral Rapid SARS-CoV-2 Antigen Rapid Test Kit and the Joysbio SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold) which use patient samples to detect proteins, called antigens, found on the SARS-CoV-2 virus. The samples are collected using a nasal swab." | "Woodside Acquisitions Inc. is recalling these tests because they were distributed to U.S. customers without authorization, clearance, or approval from the FDA. As these tests were not authorized, cleared, or approved by FDA, there is not sufficient data demonstrating that the test’s performance is accurate. This means there is a risk of potential false negative, false positive, or misinterpretation of results. If you use the affected product, this may cause serious adverse health consequences or death. Woodside Acquisitions Inc. has received no complaints or reports of injuries, deaths, or adverse events."
FDA MD Recalls | HVAD Pump Implant Kit | "The HeartWare Ventricular Assist Device (HVAD) System is used to help the heart continue to pump blood to the rest of the body. The HVAD system is used as a bridge to cardiac transplants in patients who are at risk of death from end-stage left ventricular heart failure, for heart tissue recovery, or as destination therapy (DT) in patients for whom heart transplants are not planned." | "Medtronic is recalling the HVAD Pump Implant Kit due to a pump weld defect. Following an inspection of explanted pumps returned to Medtronic, an analysis showed moisture had entered the center post of the pump causing corrosion and demagnetization of the internal magnets, which may cause the pump to rotate incorrectly. Patients with affected devices may present with signs and symptoms that resemble pump thrombosis. If this issue occurs, it may lead to pump malfunction, death, or severe injury (for example: shock with severe organ dysfunction, stroke), or require patients to undergo major surgery to replace the pump. Medtronic has received three complaints regarding this device issue, including one death, and two injuries."
FDA MD Recalls | Philips Respironics V60 and V60 Plus Ventilator | "The Philips Respironics V60 and V60 Plus ventilators are intended to support patient breathing. They are used to provide mechanical ventilation for adults and children who can breathe on their own, for example those with respiratory failure, chronic respiratory insufficiency, or obstructive sleep apnea. The ventilators are used in hospitals or other institutional settings under the direction of health care professionals." | "Philips Respironics is recalling all V60 and V60 Plus ventilators because an internal power fluctuation may cause the ventilator to shut down unintentionally either with or without any visible or audible alarms. The power fluctuation causes a reboot of the backup alarm controller, which may lead to a complete shutdown of the ventilator without warning. Affected ventilators may cease ventilation either with or without alarms. Such failures may lead to patients being deprived of oxygen for an extended time, which could cause serious adverse health consequences and death. As of April 14, 2022, there have been reports of 4 injuries and 1 death associated with the use of the recalled devices."
FDA MD Recalls | iCast Covered Stent | "The iCast Covered Stent System is a balloon expandable stent that can be delivered using a catheter to hold open and support the walls of the structures within the body. It is currently indicated for use in the United States as a stent to hold open the airways (tracheobronchial stent)." | "Atrium Medical Corporation is recalling the iCast Covered Stent System after receiving increased customer complaints about the separation of the balloon or catheter hub from the delivery system when the delivery system is being removed from a person. This issue seems to occur most often when the stent system is used outside of the indications for use, for example to treat vascular conditions. If a separation occurs and the balloon is not fully deflated before the device is removed from a person’s blood vessel, it may cause the procedure to take longer than expected, exposing the person to additional anesthesia and/or imaging chemicals known as contrast. The separated portion also may remain in the body after the delivery catheter is removed, potentially leading to vascular occlusion. Depending on the location of failure, occlusion may lead to amputation, embolism, loss of organ function, organ infarction, and tissue infarction. There have been 75 complaints, nine injuries, and no deaths related to this device issue."
FDA MD Recalls | ArjoHuntleigh Polska Sara Plus floor lift | "The ArjoHuntleigh Polska Sara Plus is an active floor lift used for short transfers, such as raising patients from bed to a wheelchair or from a wheelchair to a toilet. It is intended for use in hospitals, nursing homes, and other health care facilities." | "ArjoHuntleigh Polska is recalling the Sara Plus floor lift following several complaints of smoke and/or flames coming out of the lift. When the battery is low, the device’s printed circuit board may overheat, leading to smoke or fire. If this situation occurs, anyone using the lift or near it could be injured, including smoke inhalation and/or burns. There have been 44 complaints about this issue. No injuries or deaths have been associated with the use of this device."
FDA MD Recalls | Dragonfly OpStar Imaging Catheter | "The Dragonfly OpStar imaging catheter with optical coherence tomography (OCT) imaging system is designed to provide imaging of blood vessels that carry blood and oxygen to the heart (coronary arteries) in people who are candidates for catheter-based, minimally invasive, interventional procedures to address coronary artery disease." | "Abbott is recalling certain lots of the Dragonfly OpStar imaging catheter because the marker band farthest from the catheter tip (proximal marker) may become loose and, in two instances, has been observed to separate from the catheter while being used on a person. A loose marker band that has separated from the device may remain in the body after the catheter is removed, potentially leading to vascular injuries, including but not limited to embolism (blockage of the vessel), thrombosis (blood clot), dissection (tear), ischemia (inadequate blood supply to the heart), infarction (heart attack), infection, or death. There have been 5 incidents and 1 injury related to this device issue. No deaths have been associated with the use of this device due to this issue."
FDA MD Recalls | Avanos Medical Cortrak*2 Enteral Access System | "The Avanos Medical CORTRAK* 2 Enteral Access System is designed to help trained health care personnel place medical feeding tubes into the stomach or small bowel of patients who need to receive nutrition through the tube." | "Avanos Medical is recalling the Cortrak*2 Enteral Access System because there have been reports of injuries and patient deaths after nasoenteric or nasogastric tubes have been misplaced when this device is used to help with their placement. This recall is being used to make updates to the device's labeling, including the instructions for use and intended uses. The updates instruct users to confirm tube placement based on their institution's protocols before using the tube to deliver nutrition. If a nasogastric or nasoenteric tube is inserted incorrectly, patients could experience damage to the vocal cords, lungs, or trachea, all of which can lead to serious injury or death. According to Avanos Medical's recall communication, there have been 60 injuries and 23 patient deaths related to misplacement of nasogastric feeding tubes while using the CORTRAK* 2 Enteral Access System, since 2015. Adverse events reported included respiratory failure, pneumothorax (collapsed lung), perforation (a hole in the wall of the lung, esophagus, or bowel), pneumonia (a lung infection), and pleural effusion (excess fluid in the space between the lungs and the chest cavity)."
FDA MD Recalls | Skippack Medical Lab SARS-CoV-2 Antigen Rapid Test (Colloidal Gold), may also be called Skippack Medical Lab SARS-CoV-2 Antigen Rapid Test Kit, Lateral Flow COVID 19 Rapid Antigen Test, SKIPPACK COVID-19 Antigen Home Test | "The Skippack Medical Lab SARS-CoV-2 Antigen Rapid Test (Colloidal Gold) uses a nasal swab sample to detect proteins, called antigens, found on the SARS-CoV-2 virus." | "SML Distribution LLC is recalling these tests because these tests were distributed to U.S. customers without authorization, clearance or approval from the FDA. In addition, SML Distribution LLC did not provide the FDA with adequate validation data to show that the test’s performance is accurate. This means there is a risk of potential false negative, false positive, or misinterpretation of results. If you use the affected product, this may cause serious adverse health consequences or death. SML Distribution has received no complaints or reports of injuries, deaths, or adverse events."
FDA MD Recalls | Accula SARS-CoV-2 Test | "The Accula SARS-CoV-2 Test is a polymerase chain reaction (PCR) test intended to detect the presence of SARS-CoV-2, the virus that causes COVID-19, based on a nasal swab sample from patients. The sample is collected by healthcare providers or by people swabbing themselves under the guidance of a healthcare provider. This test is authorized for use in point of care (POC) settings. The sample is processed using special PCR laboratory equipment." | "Mesa Biotech is recalling the Accula SARS-CoV-2 Test because certain lots of the test have an increased risk of giving false positive results due to contamination at the manufacturing facility. Although there have been no reports of injuries, adverse health consequences, or death associated with the use of these affected products, false positive results could lead to further exposure of uninfected individuals to SARS-CoV-2 virus."
FDA MD Recalls | Celltrion DiaTrust COVID-19 Ag Rapid Test | "The Point of Care (POC) Celltrion DiaTrust COVID-19 Ag Rapid Test is intended to detect the presence of SARS-CoV-2, the virus that causes COVID-19, based on a nasopharyngeal (deep inside the nose to the back of the throat) swab sample from patients. The sample is collected by healthcare providers." | "Celltrion USA is recalling the affected POC DiaTrust COVID-19 Ag Rapid Test because it may have been distributed to unauthorized users. This test is authorized for use only by laboratories with CLIA certification and a corresponding CLIA Certificate Number (CCN). CLIA-certified laboratories must meet specific criteria that qualify them to perform diagnostic tests. Emergency use of these tests is limited to authorized laboratories and use of these tests in unauthorized settings (for example, customers without CLIA -certification) may present an increased risk of false results as the persons performing these tests may not have the qualifications or training to properly and safely administer them. Although there have been no reports of injuries, adverse health consequences, or death associated with the use of these affected products, false positive or false negative results from improper use of these tests could lead to further exposure of uninfected individuals to SARS-CoV-2 virus. There are also serious injury risks if someone who is not trained to collect a nasopharyngeal swab sample attempts to do so."
FDA MD Recalls | Harmony Delivery Catheter System | "The Harmony Transcatheter Pulmonary Valve (TPV) System is used to treat a leaky native or surgically repaired right ventricular outflow tract (RVOT), which is the part of the heart that carries blood to the lungs. The Harmony TPV System consists of a transcatheter pulmonary valve and a delivery catheter, the Harmony Delivery Catheter, that is used to implant the replacement valve without open heart surgery. The system is indicated for use in children and adults who have severe pulmonary regurgitation, which is when blood flows backward from the RVOT into the right lower chamber (right ventricle) of the heart, and require replacement of the pulmonary valve." | "Medtronic is recalling the Harmony Delivery Catheter because it is possible that the bond holding the capsule at the end of the delivery catheter may break during a procedure to place the TPV. A capsule bond break could cause procedure delays while the device is replaced with a new one or it may require the patient to undergo additional surgeries. Additionally, a capsule bond break while in use during a procedure could cause serious harm to the patient. Those risks include preventing blood flow and/or completely blocking (embolization or occlusion), tearing and/or splitting (perforation or dissection), or other types of damage to the patient’s blood vessels. There have been 6 reported complaints from clinical cases, one injury, and no deaths associated with the use of these devices."
FDA MD Recalls | Celltrion DiaTrust COVID-19 Ag Rapid Test (labeled for Research Use Only) | "According to the labeling for the affected products, the Point of Care (POC) Celltrion DiaTrust COVID-19 Ag Rapid Test is labeled for Research Use Only (RUO) and not for use in diagnostic procedures. The affected POC Celltrion DiaTrust COVID-19 Ag Rapid Test (labeled for Research Use Only) is a lateral flow immunoassay intended for the qualitative detection of nucleocapsid and receptor binding domains (RBDs) from the to detect the presence of SARS-CoV-2 spike proteins in human nasopharyngeal swab specimens. Results are for the identification of SARS-CoV-2 nucleocapsid and RBD protein antigen." | "Celltrion USA is recalling these specific lots of the POC DiaTrust COVID-19 Ag Rapid Test that are labeled for Research Use Only (RUO), because they were distributed to customers who would be unlikely to have used these tests for research use only and may not have the proper training to safely collect nasopharyngeal swab samples from people. Additionally, the labeling for the affected products includes uses that are outside of any authorized use (EUA210190) (for instance, use for someone suspected of COVID-19, testing is not limited to within 7 days of symptom onset, and for someone without symptoms, testing is not limited to repeat, or serial testing, performed twice between 24 and 48 hours after the first test). Emergency use of this test is limited to authorized laboratories. Although there have been no reports of injuries, adverse health consequences or death associated with the use of these affected products, false positive or false negative results from improper use of these tests could lead to further exposure of uninfected individuals to SARS-CoV-2 virus. There are also serious injury risks if someone who is not trained to collect a nasopharyngeal swab sample attempts to do so."
FDA MD Recalls | Philips Respironics V60 and V60 Plus Ventilator | "The Philips Respironics V60 and V60 Plus ventilators are intended to support patient breathing. They are used to provide mechanical ventilation for adults and children who can breathe on their own, for example those with respiratory failure, chronic respiratory insufficiency, or obstructive sleep apnea. The ventilators are used in hospitals or other institutional settings under the direction of health care professionals." | "Philips Respironics is recalling Certain V60 and V60 Plus ventilators because a subset of these devices had parts that were put together using an expired adhesive. If the adhesive fails, it could cause a capacitor support bracket to become loose and potentially damage the capacitors, which could cause the ventilator to stop providing ventilation to the patient. This failure may cause an alarm to notify the health care provider, or it may not sound any alarm at all. If ventilation stops and sounds an alarm, patients could be without ventilation support for the duration required for an alternate ventilator to be connected. If the ventilator does not sound an alarm when the ventilation stops, the patient could be deprived of oxygen for an extended time, which could cause serious adverse health consequences and death. There have been no reports of injuries, adverse health consequences, or death associated with the use of these devices."
FDA MD Recalls | SD Biosensor STANDARD Q COVID-19 Ag Home Test | "The SD Biosensor STANDARD Q COVID-19 Ag Home Test uses a nasal swab sample to detect proteins, called antigens, from SARS-CoV-2, the virus that causes COVID-19. This test is not authorized, cleared, or approved by the FDA for marketing or distribution in the United States. Note: This recall does not apply to the SD Biosensor COVID-19 At-Home Test (also an antigen test), which was authorized by the FDA on December 24, 2021, and is distributed by Roche Diagnostics." | "SD Biosensor Inc. is recalling its STANDARD Q COVID-19 Ag Home Tests because these tests were distributed to U.S. customers without authorization, clearance, or approval from the FDA. As this test was not authorized, cleared, or approved by the FDA, there is not sufficient data demonstrating that the test's performance is accurate. This means there is a risk of both false-negative and false-positive test results. False negative results are when the test does not detect the SARS-CoV-2 virus but the person is actually infected. False-positive results occur when the test says the person has SARS-CoV-2 virus present, but they are not infected. There have been no reports of injuries, adverse health consequences, or death associated with the use of the SD Biosensor STANDARD Q COVID-19 Ag Home Test. Use of the affected product could cause serious adverse health consequences and death. On March 1, 2022, FDA issued a Safety Communication, Do Not Use SD Biosensor STANDARD Q COVID-19 Ag Home Tests, warning users to stop using these tests."
FDA MD Recalls | Celltrion DiaTrust COVID-19 Ag Rapid Test | "The Point of Care (POC) Celltrion DiaTrust COVID-19 Ag Rapid Test is intended to detect the presence of SARS-CoV-2, the virus that causes COVID-19, based on a nasopharyngeal (deep inside the nose to the back of the throat) swab sample from patients. The sample is collected by healthcare providers." | "Celltrion USA is recalling specific lots of the POC DiaTrust COVID-19 Ag Rapid Test due to a high number of false positive reports. A false-positive antigen test result may lead to a delay in both the correct diagnosis and treatment for the actual cause of a person's illness. False-positive results could also lead to more spread of the SARS-CoV-2 virus if presumed positive people are housed together. Additionally, the tests’ labeling for the affected POC products includes a shelf life of 18 months. However, the FDA’s emergency use authorization specifies these tests may only be used for 12 months. The use of the affected product could cause serious adverse health consequences and death. Although there have been no reports of injuries, adverse health consequences or death associated with the use of the affected POC products, false positive or false negative results from improper use of these tests could lead to further exposure of uninfected individuals to SARS-CoV-2 virus. There are also serious injury risks if someone who is not trained to collect a nasopharyngeal swab sample attempts to do so."
FDA MD Recalls | COVID-19 Antigen Test (Saliva) COVID-19 Antigen Test (Nasal) COVID-19 IgG/IgM Antibody Test | "The LuSys Laboratories COVID-19 Antigen Test uses a nasal swab, or a saliva (spit) sample intended to detect proteins, called antigens, on the SARS-CoV-2 virus. The LuSys Laboratories COVID-19 IgG/IgM Antibody Test uses serum, plasma, or blood samples to look for antibodies produced by a person's immune system in response to SARS-CoV-2, the virus that causes COVID-19, suggesting a recent or previous infection. Antibody tests should not be used to diagnose or exclude an active COVID-19 infection. The LuSys Laboratories COVID-19 Antigen Tests (Nasal/Saliva) and COVID-19 IgG/IgM Antibody Tests may also be sold under the company names Luscient Diagnostics, Vivera Pharmaceuticals, or with the trade name EagleDx." | "LuSys Laboratories is recalling these tests because they do not have an Emergency Use Authorization, 510(k), or PMA and therefore cannot be legally marketed and distributed in the United States. In addition, LuSys Laboratories did not provide appropriate validation data to show that the tests can perform accurately. This means there is a risk of potential false negative, false positive, and misinterpretation of results from these tests. LuSys Laboratories has received no complaints or reports of injuries, deaths, or adverse events."
FDA MD Recalls | Baxter SIGMA Spectrum Infusion Pump with Master Drug Library (Version 8) Spectrum IQ Infusion System with Dose IQ Safety Software (Version 9) | "The SIGMA Spectrum Infusion Pump with Master Drug Library (Version 8) and Spectrum IQ Infusion System with Dose IQ Safety Software (Version 9) are software-controlled infusion pumps that deliver controlled amounts of fluids such as pharmaceutical drugs, blood, and blood products and other required patient therapies. The fluids are provided through an infusion tubing set into a patient's vein or through other cleared routes of administration. The devices are used in hospitals and other healthcare facilities." | "Baxter Healthcare is recalling SIGMA Spectrum Infusion Pumps with Master Drug Library (Version 8) and Spectrum IQ Infusion Systems with Dose IQ Safety Software (Version 9) due to the risk of not alarming for repeated upstream occlusion events. After an upstream occlusion alarm, failure to fully resolve any upstream occlusion before restarting the infusion can cause the pump to not re-alarm as expected. The pump adjusts occlusion detection after the user acknowledges an upstream occlusion alarm; for the remainder of that infusion, the upstream occlusion alarm will not alarm per its baseline settings. If an occlusion is not fully resolved before restarting the infusion, the pump may not re-alarm as expected and may appear to be infusing normally, though may be infusing below the programmed rate or not infusing at all. Patient harm may occur from interruption in therapy (due to full occlusion) and/or underinfusion (due to partial occlusion). Use of the affected products may cause serious adverse events, including death. Baxter has received 51 reports of serious injuries and three reports of patient deaths over five years, potentially associated with this issue."
FDA MD Recalls | TurboHawk Plus Directional Atherectomy System | "The TurboHawk Plus Directional Atherectomy System consists of a catheter and a cutter driver. This device is used during procedures intended to remove blockage from peripheral arteries and improve blood flow." | "Medtronic is recalling (correcting) this product due to design similarities shared with another device that was recently recalled for correction. There is a risk of the guidewire within the catheter moving downward or prolapsing (See Figure 1) when force is applied during use. If this happens, the catheter tip may break off or separate and this could lead to serious adverse events including a tear along the inside wall of an artery (arterial dissection), a rupture or breakage of an artery (arterial rupture), decrease in blood flow to a part of the body because of a blocked artery (ischemia), and/or blood vessel complications that could require surgical repair and additional procedures to capture and remove the detached and/or migrated (embolized) tip. As of February 7, 2022, there have been no reported injuries or deaths associated with this problem."
FDA MD Recalls | Arrow-Trerotola Percutaneous Thrombolytic Device (PTD) | "The Arrow-Trerotola Percutaneous Thrombolytic Device (PTD) consists of a catheter with an expandable basket attached to a drive cable. The catheter is used in with a low-speed, battery-operated rotator unit to remove blood clots in native arterio-venous fistulae and synthetic dialysis grafts." | "Arrow-International, LLC (a subsidiary of Teleflex Inc.) is recalling the Arrow-Trerotola Percutaneous Thrombolytic Device due to risk of tip damage during use, which may result in tip detachment from the basket. This could potentially lead to vascular injuries, including obstruction (blockage) of the vessel, additional thrombosis (blood clot), ischemia (inadequate blood supply), infarction (heart attack), infection or death. As of December 2021, a total of 35 complaints reporting tip separation have been received. Of these 35 complaints, 14 reported injuries, and nine complaints involved use of a stent (support structure) to manage the separated tip. No deaths have been reported at this time."
FDA MD Recalls | E25Bio COVID-19 Direct Antigen Rapid Test; Trade Name: E25Bio SARS-CoV-2 Antigen Test Kit | "The E25Bio COVID-19 Direct Antigen Rapid Test is used to detect proteins called antigens from the SARS-CoV-2 virus in patient samples. These samples are collected in one of three ways: Nasal (anterior nares) swab Nasopharyngeal swab, which reaches deep inside the nose, reaching the back of the throat or from the middle part of the throat (pharynx) Oropharyngeal swab reaching just beyond the mouth into the throat" | "E25Bio is recalling its COVID-19 Direct Antigen Response Tests (DART) for several reasons, particularly that these tests were marketed and distributed to U.S. customers without authorization, clearance, or approval from the FDA. Labeling distributed with some of the tests also includes inaccurate claims and instructions, including a statement that misrepresents the test as FDA-authorized. As this test was not authorized, cleared, or approved by the FDA, there is not sufficient data demonstrating that the test’s performance is accurate. This means there is a risk of both false-negative and false-positive test results. False negative results are when the test does not detect the SARS-CoV-2 virus but the person is actually infected. False-positive results occur when the test says the person has SARS-CoV-2 virus present but they are actually not infected. There is also a risk of injury if users follow any labeling instructions directing self-collection of nasopharyngeal or oropharyngeal samples. Only trained health care providers should collect these types of swab samples to prevent serious injury. Use of the affected product could cause serious adverse health consequences and death. On February 4, 2022, the FDA issued a Safety Communication warning users to stop using these tests. There have been no reports of injuries, adverse health consequences or death associated with the use of this product."
FDA MD Recalls | RNAstill MTM specimen collection kits | "BASE10 Genetics RNAstill MTM specimen collection kits are intended for collection, transport, inactivation, stabilization, and long-term storage of nasopharyngeal (deep inside the nose to the back of the throat) or oropharyngeal (just beyond the mouth into the throat) swab samples from patients, without needing dry ice, refrigeration, or special containment. They can be used to collect and store patient samples that need to be tested for viruses including SARS-CoV-2 (the virus that causes COVID-19) and other coronaviruses, influenza (flu), or adenoviruses (viruses that typically cause the common cold). They can also be used for patient samples that need to be tested for bacteria that commonly cause infections including Mycoplasma pneumoniae, Chlamydia pneumoniae, Streptococcus pneumoniae, and Streptococcus pyogenes." | "BASE10 Genetics is recalling the RNAstill MTM specimen collection kits because they were distributed without proper premarket clearance from the FDA—meaning there is not adequate data available about this product to know how well it inactivates and stabilizes a virus for transport and storage. If the specimen collected is not properly inactivated and stabilized by the transport medium, there are two concerns: If the sample deteriorates, it could lead to false negative test results If the virus isn’t inactivated prior to testing, it is possible that the virus could spread to people in the lab. These products were distributed to nursing homes, assisted living facilities, and memory facilities whose residents are at high risk of severe illness and death from COVID-19. A false negative test result could lead to a patient not receiving the care they need. Additionally, the transport medium of this specimen collection kit contains a hazardous chemical that requires special training for safe handling. It is not certain that staff at long term care facilities where the product was distributed are trained to safely handle this hazardous chemical. As a result of all these issues, use of the affected product could cause serious adverse health consequences and death. There have been no reports of injuries, adverse health consequences or death associated with the use of this product."
FDA MD Recalls | Vyaire Medical Bellavista 1000 and 1000e Series Ventilators | "The bellavista 1000 and 1000e Series Ventilators are used to assist patient breathing by delivering air, or air with extra oxygen, into the airways or lungs. The ventilators can be used on adult and pediatric patient populations, and optionally, on infant and neonatal patients. When a patient needs oxygen, these ventilators can be used at bedside in hospitals, intensive care (sub-acute) facilities, and can be transferred within a facility." | "Vyaire Medical is recalling the bellavista 1000 and 1000e Series Ventilators, with specific software versions and configurations, after reports of ventilators failing to ventilate and generating a technical failure alarm. Issues with the software version 6.0.1600.0 or higher installed can have a conflict in memory between software tasks when the data communication port is set to "HL7," which produces the technical failure alarm 305. The use of affected ventilators may cause the ventilator to malfunction or stop, which may cause serious adverse events. There have been 18 complaints, seven injuries, and no reports of death."
FDA MD Recalls | Empowered Diagnostics, LLC CovClear COVID-19 Rapid Antigen Test and ImmunoPass COVID-19 Neutralizing Antibody Rapid Test | "The CovClear COVID-19 Rapid Antigen Test uses a nasal swab sample to detect proteins, called antigens, from SARS-CoV-2, the virus that causes COVID-19. The ImmunoPass COVID-19 Neutralizing Antibody Rapid Test uses a fingerstick blood sample to detect antibodies produced by a person’s immune system in response to SARS-CoV-2. If antibodies are found it means the person may previously have been infected with the SARS-CoV-2 virus. Antibody tests should not be used to diagnose or exclude an active COVID-19 infection." | "Empowered Diagnostics is recalling the CovClear COVID-19 Rapid Antigen Test and ImmunoPass COVID-19 Neutralizing Antibody Rapid Test. These tests were distributed with labeling indicating they are authorized by the FDA, but neither test has been authorized, cleared, or approved for distribution or use in the United States by the FDA. The FDA is concerned about the potentially higher risk of false results when using unauthorized tests." There have been no reports of adverse health consequences or death from use of these tests."
FDA MD Recalls | Trilogy Evo ventilator Repair kits for Trilogy Evo muffler assembly* | "The Trilogy Evo ventilator provides continuous or intermittent positive pressure ventilation for the care of pediatric through adult patients weighing at least 2.5 kg who require mechanical ventilation. The ventilator can measure, display, record, and alarm oxygen saturation level (SpO2), fraction of inspired oxygen (FiO2), carbon dioxide (CO2) and pulse rate data when integrated with the appropriate accessories. The ventilator is suitable for use in institutional, home, and non-emergency transport settings, for example wheelchair or personal vehicle. It may be used for both invasive and non-invasive ventilation." | "Philips Respironics has recalled certain Trilogy Evo ventilators with specific serial numbers due to potential health risks. A Philips supplier incorrectly used polyester-based polyurethane (PE-PUR) sound abatement foam, a non-conforming material, in the muffler assembly of the affected Trilogy Evo ventilators. The issue was identified during lab testing of the Trilogy Evo ventilator. The Trilogy Evo ventilators with non-conforming foam were distributed to customers in the United States and Korea. There have been no reported injuries or death to date. The polyester-based polyurethane (PE-PUR) sound abatement foam, which is used to reduce sound and vibration in these affected devices, may break down and potentially enter the device's air pathway. If this occurs, black debris from the foam or certain chemicals released into the device's air pathway may be inhaled or swallowed by the person using the device. For more information, please see the Potential Health Risks from Sound Abatement Foam section of the Update: Certain Philips Respironics Ventilators, BiPAP, and CPAP Machines Recalled Due to Potential Health Risks: FDA Safety Communication."
FDA MD Recalls | HawkOne Directional Atherectomy System | "The HawkOne Directional Atherectomy System consists of a catheter and a cutter driver. This device is used during procedures intended to remove blockage from peripheral arteries and improve blood flow." | "Medtronic Inc. is recalling this product due to the risk of the guidewire within the catheter moving downward or prolapsing when force is applied during use (see Figure 1). If this happens, the catheter tip may break off or separate and this could lead to serious adverse events including a tear along the inside wall of an artery (arterial dissection), a rupture or breakage of an artery (arterial rupture), decrease in blood flow to a part of the body because of a blocked artery (ischemia), and/or blood vessel complications that could require surgical repair and additional procedures to capture and remove the detached and/or migrated (embolized) tip. There have been 163 complaints about this device issue. There have been 55 injuries and no deaths reported about this device issue."
FDA MD Recalls | Vaporizer Sevoflurane, Maquet Filling for Flow Family Anesthesia Systems | "The Vaporizer is a component of several Maquet anesthesia gas machines. The Flow Family of Anesthesia System administers inhalation anesthesia while controlling the entire ventilation of patients who are unable to breathe, as well as in supporting patients with a limited ability to breathe." | "Getinge is recalling the Vaporizer Sevoflurane Maquet filling due to the potential chemical breakdown of Sevoflurane, a general surgical anesthetic, that may result in inhalation and/or skin exposure to harmful chemicals. If this occurs, this may cause serious patient harms including irritation of the respiratory tract, lung edema (swelling caused by excess fluid), and severe hypocalcemia (condition in which the blood has too little calcium). There have been eight complaints regarding this device issue. There have been no reported deaths or injuries."
FDA MD Recalls | WIRION Embolic Protection Device | "The WIRION® system is used to hold and remove debris or blood clots from the lower limbs that can be associated with atherectomyExternal Link Disclaimer procedures." | "Cardiovascular Systems Inc. is recalling the WIRION system due to complaints of filter breakage during retrieval. Under certain circumstances, such as when the filter basket is too full, the filter assembly may become difficult to withdraw. In this situation, withdrawal may cause the WIRION system filter component to tear or separate, which may result in series adverse events such as embolizationExternal Link Disclaimer, need for additional medical procedures, or possibly death. Cardiovascular Systems Inc. has received reports of nine device malfunctions and no reports of death related to this device issue."
FDA MD Recalls | Synergy Cranial and StealthStation S7 Cranial Software | "The StealthStation System, with Synergy Cranial software or StealthStation S7 Cranial software, aids in precisely locating anatomical structures in neurosurgical procedures." | "Medtronic is recalling its Synergy Cranial and StealthStation S7 Cranial software due to the potential inaccuracies caused by the Biopsy Depth Gauge Cycle View. If the user encounters the software issue where the graphical Biopsy Depth Gauge is no longer synchronized with other navigation views, it may result in a prolonged procedure, the need for an additional surgical procedure, aborted procedure, tissue injury, including potential for life-threatening injury (such as hemorrhage, unintended tissue damage, or permanent neurological injury), which could lead to death. There have been four complaints regarding this device issue. There have been no reported injuries or deaths."
FDA MD Recalls | Puritan Bennett 980 Series Ventilator | "Ventilators are machines that provide respiratory support by acting as bellows to move air in and out of your lungs. Your health care provider sets the ventilator to control how often it pushes air into your lungs and how much air you get." | "Covidien, LP (part of Medtronic) is recalling its Puritan Bennett 980 Series Ventilator due to a capacitor manufacturing assembly error that may cause the ventilator to become inoperable or stop working as intended. If this occurs, it could result in the loss of ventilation and serious adverse events such as hypercarbia (an increase of carbon dioxide in the blood), hypoxemia (an abnormally low amount of oxygen in the bloodstream), neurological injury or death. There have been six complaints and one death regarding this device issue. There have been no reported injuries."

## FDA MDR 2023
Base | Device name | Device use | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
FDA MD Recalls | Panorama 1.0T HFO | "The Panorama 1.0T is a magnetic resonance (MR) system, a machine used to take different kinds of images. These images show what's inside the body and can help health care providers make a diagnosis." | "Philips North America LLC is recalling the Panorama 1.0T HFO due to risk of explosion during a quench procedure caused by excessive pressure buildup of helium gas. During a quench, which is not common, a large amount of helium evaporates and is vented outside the building through a venting system. If an unknown blockage is present in the venting system and the pressure exceeds design limits, the structural integrity of the system could be compromised. The magnet in the Panorama 1.0T HFO may experience an unintended quench during normal use or when initiated in an emergency by the operator pressing the Magnet EMERGENCY STOP button. The use of an affected system may cause serious adverse health consequences, including chemical exposure, lack of oxygen, tissue damage, and mechanical trauma caused by debris, such as brain injury, laceration, fracture, eye injury, bruises, and death. The issue could also lead to system or property damage. There has been one reported event of an explosion in 22 years of use. There have been no reports of injury or death."
FDA MD Recalls | Olympus bronchofiberscopes and bronchovideoscopes | "The bronchofiberscopes and brochovideoscopes are tubular devices that work with accessories like cameras and lights to examine or treat a person’s airways." | "Olympus Corporation of the Americas (Olympus) is recalling certain bronchofiberscopes and bronchovideoscopes because they may lead to burns or catch fire. Using the affected bronchoscopes may cause serious adverse health consequences, including critical burns in a person’s airways or lungs, airway bleeding, trouble breathing, apnea, loss of consciousness, or death. Injuries may lead to prolonged procedures, additional medical care, extended hospitalization, ICU care, and death. Combustion can damage or break parts of the device that may injure or unintentionally remain in the patient and may require retrieval or surgical removal. There have been 192 complaints regarding this issue, including four injuries. There have been no deaths reported. There is a risk of combustion if: high-frequency cauterization is performed while supplying oxygen, or the electrode section of an electrosurgical accessory is too close to the end of the endoscope."
FDA MD Recalls | StealthStation S8 Application Version 2.0 and 2.0.1 | "The StealthStation System is used for any medical condition in which the use of stereotactic surgery may be appropriate. This system keeps track of where surgical tools are in relation to a patient's anatomy by using images of the patient. As the surgeon works, the system shows where the tool is on these images and models where the surgeon picked. Before surgery, the surgeon can plan the path for the tools and see how they should move. During surgery, the software helps the surgeon see if the actual tool position matches the plan, guiding them along the right path. These tools help find body parts during brain surgery, whether it's open or less invasive. They're useful for any medical situation where precise surgery might be needed." | "Medtronic Navigation Inc. is recalling StealthStation S8 Application Version 2.0 and 2.0.1 due to a software glitch that can make the surgical plan data shift location(s) after the initial test is changed. This can cause the surgical plan data to be in an unintended location for the surgery. The use of affected product may cause serious adverse health consequences, including death. There have been no reported injuries. There have been no reports of death."
FDA MD Recalls | CARDIOHELP Emergency Drive | "The CARDIOHELP System helps oxygenate blood and remove carbon dioxide from it. It moves blood through a circuit outside the body to support the heart and lungs for up to six hours. It is also intended to support the heart and lungs during procedures not requiring cardiopulmonary bypass, again for up to six hours." | "Getinge is recalling the CARDIOHELP Emergency Drive due to a possible blocking or impairment of the ability of the emergency drive to be cranked. The emergency drive can become stuck or difficult to turn due to friction being generated in the handle attachment. The stiff hand crank is caused by a scrounger disc that interferes with the turning of the crank. If the emergency drive is needed, a clinician may not be able to turn the handle to drive the pump or may not be able to turn the handle as fast as is needed to adequately support the patient. If this occurs, the patient will lose adequate hemodynamic support or gas exchange; the results of which could include ischemia, hypoxia, stroke, or death. There have been no reports of death or injuries associated with this issue."
FDA MD Recalls | Cordis INFINITI Angiographic Catheter | "The Cordis INFINITI Angiographic Catheter delivers contrast dye to specific places in the blood vessels. It helps doctors see the blood vessels in the brain, organs, or limbs during certain procedures. The catheter is put into the body through the skin and can be used in multiple places at once." | "Cordis US Corp is recalling INFINITI Angiographic Catheter due to some of the products being shipped to end users without undergoing sterilization procedures. The INFINITI Angiographic Catheter is intended to be shipped sterile directly from the manufacturer to a third-party distributor, which in turn ships the product to end users such as hospitals. Lot 18223410 was distributed without being properly sterilized. All units, except 3 boxes (30 units of the 500 units in the lot), were sent back to the manufacturer in Mexico. A review of shipments revealed that these 30 units had been shipped to a single facility. Cordis US Corp also phoned the sole consignee on September 26, 2023 to notify them of this issue. Cordis US Corp then issued an Urgent Medical Device Recall Notice to its sole consignee on September 28, 2023 to request the return of all affected product. The use of unsterilized angiographic catheters may result in serious acquired infections, sepsis, and in extreme cases, death. There have been no reports of death or injury associated with this issue."
FDA MD Recalls | BD Alaris Syringe Module, BD Alaris PCU, and Alaris PCA Module | "The Alaris infusion pump is a modular infusion pump system generally intended to deliver medications or other fluids. The syringe and PCA pump modules are validated for use with specific syringes, as reflected in the product labeling. On the Alaris system, the user enters the nominal volume of the syringe and selects the syringe brand from a library of compatible syringes. This instructs the PCU on which pre-loaded syringe dimensions to use when calculating the volume of fluid in the syringe. Monoject is one of the most common syringes used with infusion pumps and is specifically mentioned as an example in the Alaris 8015 user manual as shown below." | "Becton Dickinson (BD)/Carefusion 303 is recalling their Alaris Infusion Pumps to update the labeling due to compatibility issues with Cardinal Health Monoject syringes. The Alaris pumps are validated for use with ‘Monoject’ syringes and list ‘Monoject’ as one of the options when the user is selecting a syringe type within the Alaris system. However, the dimensions for Monoject syringes have recently changed while rebranding the syringes from “Covidien Monoject” to “Cardinal Health Monoject.” Cardinal Health branded Monoject syringes have not been validated for use with the BD Alaris Syringe and PCA modules. When one of these new syringes is used with the previous syringe settings, the pump may either refuse to operate or incorrectly estimate the volume of liquid in the syringe. This has resulted in over and under infusions, as well as delays in therapy. Use of incompatible syringe sizes and models with the BD Alaris Syringe and PCA Modules can impact syringe pump operation resulting in inaccurate fluid delivery, delayed generation of occlusion alarms, and other potential problems. There have been 13 reported injuries. There have been no reports of death."
FDA MD Recalls | Eitan Medical Ltd Sapphire Infusion Pumps running software version Rev 16.10 | "The Sapphire infusion pumps deliver fluids, nutrition, including lipids, as well as medication, blood, and blood products such as platelets and plasma. The pumps are designed to be used in hospitals, clinics, at home, or they may be portable. The pumps run under both battery power and when connected to a power source." | "Eitan Medical Ltd is recalling the Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, and Sapphire Plus Infusion Pump due to software issues in software version Rev 16.10. Pumps with this software version may fail to detect air in the line when running on battery power. In some cases, the pump may not sound an alarm when there is air in the line. Failure to detect air in the line can lead to blockage in a blood vessel (embolism), causing complications such as unstable blood pressure, stroke, heart attack, or even death. There have been no reports of injury or death associated with this issue."
FDA MD Recalls | VariSoft Infusion Sets | "VariSoft is an infusion set used with Tandem insulin pumps for the subcutaneous infusion of insulin in the treatment of diabetes. With an insertion angle between 20-45°, the infusion set allows users to choose their most comfortable insertion angle. It is often preferred by users who are thin or who have scar tissue or limiting potential insertion sites. This product is delivered sterile and ready for use. " | "Unomedical A/S is recalling certain VariSoft Infusion Sets because there is a risk the connector may detach from the infusion set which will disrupt the delivery of the insulin to the patient. If this occurs, insulin may continue to flow through the tubing that is now disconnected from the patient and the insulin pump would not recognize an infusion set problem and would not alert users. The defect reported was caused by damage to the connector piece during manufacturing, resulting in the connector becoming detached more easily from the insulin set than expected. Disconnections occurring during sleep, or when otherwise not detected, may lead to hyperglycemia unless it is noticed and reconnected promptly. There is a chance of death or life-threatening harm to individual patients using devices from these lots, requiring professional medical intervention. There has been 1 reported injury. There have been no reports of death."
FDA MD Recalls | Infusomat Space Large Volume Pump, Wireless and Infusomat Space Large Volume Pump, Non-Wireless BATTERY PACK | "The Infusomat Space Volumetric Infusion Pump System, which includes the Infusomat Space Large Volume Pump, Wireless and Infusomat Space Large Volume Pump, and Non-Wireless BATTERY PACK, is intended to be used by trained healthcare professionals in healthcare facilities for use on adults, pediatrics, and newborns for the intermittent or continuous delivery of parenteral fluids, medications, blood, and blood products." | "B. Braun Medical, Inc. is recalling the Infusomat Space Volumetric Infusion Pump System due to faulty occlusion alarms. On certain models, an occlusion alarm may sound when no occlusion exists, causing the pump to stop the delivery of medications, such as high-risk medications like vasopressors. Interruption to the infusion of high-risk medications may lead to hemodynamic instability requiring medical intervention to prevent permanent impairment to body structures or body functions. In some cases, this may be life-threatening or may lead to death. There have been 51 complaints, one reported injury, and one death related to this recall."
FDA MD Recalls | Asensus Senhance Surgical System | "The Senhance Surgical System is intended for use in general laparoscopic surgical procedures and laparoscopic gynecological surgery. It is intended for use by trained physicians in an operating room environment in accordance with the instructions for use. Use of the device is limited to patients two (2) years of age and older and a weight equal to or above 10kg, who are suitable to be subjected to a conventional endoscopic technique." | "Asensus Surgical Inc. is recalling the Senhance Surgical System due to malfunctions that cause unintended movement of the robotically-assisted surgical device. This issue presented itself as an uncontrolled arm motion of the Laparoscope Instrument Actuator (LIA) where the LIA rotated continuously in one direction after the Surgeon removed the engagement of teleoperation on the Senhance System. The Senhance System is designed with emergency stop capabilities to stop this issue if observed. There has been no patient impact or harm that has occurred due to this issue, however, the potential for critical tissue trauma is possible."
FDA MD Recalls | Sanxin Single Use Sterile Syringes | "The Fresenius Medical Care Sanxin Single Use Sterile Syringes are used during hemodialysis treatments. They are used to draw heparin, a blood thinning medicine used to prevent blood clots. The heparin is then mixed with the blood in the hemodialysis machine using a mechanical syringe pump. The syringes are connected to the Fresenius 2008T hemodialysis machine and pressurized for the duration of the hemodialysis treatment." | "Fresenius Medical Care is recalling its Sanxin Single Use Sterile Syringes due to blood or heparin leaking back or from the syringe. There have also been reports of an unknown black material inside the syringe. The use of affected syringes may cause serious adverse health consequences, including blood stream infection (sepsis), or blood loss due to leakage. Other risks include the wrong heparin dosage, leading to blood clotting, embolism, and death. There have been 37 reported incidents and no reported injuries or deaths."
FDA MD Recalls | Novum IQ Syringe Pump | "The Novum IQ syringe pump is an infusion pump, intended to help deliver an accurate amount of fluids, whether nutrients or medications, into a patient's body in a controlled manner. The device is suitable for patient care in hospitals and outpatient healthcare facilities for use on adults, pediatrics, and neonates by trained health care professionals." | "Baxter Healthcare Corporation is recalling the Novum IQ Syringe Infusion Pump because the pump may indicate an infusion is complete when it is not. The problem has been traced to a software error that may miscalculate volume after the pump detects a blockage (occlusion). Use of the pump could lead to serious adverse health consequences and death. People treated with the infusion pump might not get enough of the fluids they need, or time-sensitive treatment could be delayed. Underdosing or delays could lead to serious adverse health consequences, especially for people receiving life-sustaining medications. In high-risk populations, insufficient fluid delivery could lead to death. There have been no reports of serious injury or death associated with the software problem."
FDA MD Recalls | Cardinal Health Monoject single use Luer lock syringes 1, 6, 12, 20, 35 and 60mL | "Cardinal Health Monoject disposable syringes are used to inject fluid into or withdraw fluids from the body. When used with syringe pumps, the Monoject disposable syringes are loaded with fluid or medications and placed into the pump. Syringe pumps deliver solutions such as fluids, medications, and blood products to patients." | "In June 2023, Cardinal Health began distributing Monoject syringes branded as “Cardinal Health Monoject syringes.” These new syringes differ from the previously branded “Covidien Monoject syringes” as they have different dimensions. The affected Cardinal Health Monoject syringes (seen on the right in the figure above) should not be used with syringe pumps. The dimensional changes made to the affected Cardinal Health Monoject syringes when used with syringe pumps may result in pump performance issues such as overdose, underdose, delay in therapy, and delays in occlusion alarms. Cardinal Health has received 15 reports of delayed therapy due to syringe infusion pumps not recognizing syringes, and 13 reports of inaccurate volume/rate dispensing, including some injuries. Cardinal Health has not received any reports of patient death."
FDA MD Recalls | McGRATH™ MAC Video Laryngoscopes | "The McGRATH™ MAC video laryngoscope is a tool used by trained and licensed individuals to gain a view of the vocal cords during medical procedures, typically a laryngoscopy. The device incorporates a light source (LED) and miniature camera (camera) to view the larynx during medical procedures." | "Covidien LLC is recalling McGRATH MAC Video Laryngoscope in order to remove stolen defective models that have been offered for sale illegally by unauthorized third parties. These Laryngoscopes, meant for destruction, were being sold on social media platforms. After the theft was discovered, some devices were retrieved, and a for sale posting on FaceBook Marketplace was taken down. However, it is unknown where the devices were distributed. There have been no complaints regarding this devices issue. The firm indicates no injuries and no deaths have been reported for this issue. All the defects and failure modes are not known to the firm because they were stolen from the intended scrap. As such, it is not possible to rule out the major failure modes that may lead to inadequate airway support, cardiopulmonary arrest, and death. Subpopulations at higher risk include those with higher metabolic rates and lower pulmonary reserve – pediatrics, pregnant, obese, critically ill."
FDA MD Recalls | Pressure Injectable Central Venous Catheters | "The Teleflex and Arrow International Pressure Injectable Catheter Kits allow health care providers access to a patient’s central cardiovascular system, which connects the heart and veins. The kits may be used short-term (less than 30 days) for patients with conditions that require frequent injections or blood draws. With the kit, a health care provider can sample blood, monitor blood pressure, or administer fluids, medications, or chemotherapy." | "Teleflex, and its subsidiary Arrow International, are recalling the below listed Pressure Injectable Catheter Kits due to mislabeling regarding the presence of chlorhexidine in these products. The product code and product name are incorrectly listed as non-coated on the lids of the affected Pressure Injectable Catheter Kits. However, the banner card correctly lists the product code and product name as chlorohexidine coated. People who inadvertently use these kits may have serious adverse health effects, such as allergic reactions including itching, redness, skin changes, hives, dizziness, drop in blood pressure, difficulty breathing, and anaphylaxis which could lead to death if they are allergic to chlorohexidine. There have been 16 reported incidents associated with this issue. There have been no reports of injuries or deaths related to this issue."
FDA MD Recalls | Olympus High Flow Insufflation Unit | "The Olympus High Flow Insufflation Unit inflates the abdomen or colon with carbon dioxide gas. The unit may be used during laparoscopic or endoscopic surgeries, such as minimally invasive abdominal surgeries, colonoscopies, or vein harvesting. The unit also provides automatic suction and smoke evacuation." | "Olympus is recalling its Olympus High Flow Insufflation Unit because the unit may over-insufflate (inflate) air into the body with no warning or alarm. Olympus Recalls Olympus High Flow Insufflation Unit Due to Over-Inflation The use of the affected unit may cause serious adverse health consequences, including air embolism, and arrythmias including bradycardia, asystole, or cardiac arrest. Other risks include collapsed lung, kidney or urinary problems, oxygen cut off to organs, and air trapped under the skin. Other necessary or complex medical procedures may be delayed because of over-insufflation, and death may occur. There have been reports of 21 malfunctions, 10 serious injuries, and one death."
FDA MD Recalls | 2008 Series Hemodialysis Machines | "The Fresenius Medical Care 2008 Series Hemodialysis Machines act as an artificial kidney. They are used for hemodialysis treatment to pump blood from a person’s body to remove toxins and excess water." | "Fresenius Medical Care is recalling some of its 2008 Series Hemodialysis Machines because patients may be exposed to non-dioxin-like (NDL) polychlorinated biphenyl acids (PCBAs) that leach from some peroxide cured silicone tubing used as part of the hydraulics in the machine and dialysate lines. Data provided by the manufacturer indicates that the amount of NDL PCBAs decreases over the first month during routine clinical use of the hemodialysis machine. Fresenius Medical Care is updating affected machines manufactured between August 21, 2008 and June 6, 2022 that contain chlorinated peroxide cured silicone tubing and may not have been used or may have been in use for less than 36 days (or 486 total hours). The use of affected machines may cause serious adverse health consequences months to years after exposure, including endocrine dysfunction, liver issues, neurobehavioral changes, skin problems (acne, rashes), and male infertility. There have been no reports of death."
FDA MD Recalls | Philips Respironics V60 Ventilators and V60 Plus Ventilators | "Philips Respironics V60 Ventilators and V60 Plus Ventilators are used to provide breathing assistance to pediatric and adult patients. A ventilator is intended to mechanically control or assist patient breathing by delivering a predetermined percentage of oxygen. V60 Ventilators and V60 Plus Ventilators are intended to be used inside hospitals." | "Philips Respironics, Inc. is recalling their V60 Ventilators and V60 Plus Ventilators. The ventilators contain Power Management Printed Circuit Board Assemblies (Power Management PCBA) that were distributed by a third-party warehouse. The Power Management PCBAs do not meet ventilators standards. The V60 Ventilators and V60 Plus Ventilators are not always used in Intensive Care Units (ICUs) where patients are directly monitored. There may be greater risk for patients who are not directly monitored in other hospital settings. The use of the affected ventilators may cause serious adverse health consequences, including power failure, loss of ventilator function, and death. There have been no reports of death or injuries related to this recall."
FDA MD Recalls | Hamilton C1/T1/MR-1 Ventilators | "Hamilton ventilators are used to support adults, children, infants and newborns with breathing. Health care professionals use them during patient transport in or outside hospital settings." | "Hamilton Medical, Inc. is recalling the HAMILTON-C1, T1, MR-1 ventilators because the capacitators may leak electrolyte fluid onto the ventilator’s control board. If the control board contacts the electrolyte fluid, the control board or installed spare parts could short circuit. As a result of the short circuit, the ventilator may switch to “Ambient State.” The ‘Realtime clock failure’ will not have severe effects, but the ‘Ambient State’ failure will require immediate action. ‘Ambient State’ may require a health care professional to provide air supply to the patient with a manual resuscitating device and use another ventilator. Serious injury or death may result for patients who do not breathe on their own or do not receive immediate intervention. There have been zero reported incidents, injuries, or deaths related to this recall."
FDA MD Recalls | Hudson RCI Addipak Unit Dose Vial, 0.9% Full Normal Saline Solution | "The Hudson RCI Addipak Unit Dose Vials is a pack of single use 0.9% Full Normal Saline Solution used during treatments such as inhalation and irrigation therapy. It may be used with a non-ventilatory nebulizer to clean the lungs or for tracheal irrigation. For inhalation therapy, the saline solution restores moisture to the lungs, relieves congestion caused by colds or allergies, or dilutes bronchodilator (rescue inhaler) solutions that require dilution. For irrigation therapy, the saline solution is used for wound cleansing and flushing." | "Medline Industries is recalling the Hudson RCI Addipak Unit Dose Vial, 0.9% Full Normal Saline Solution (Lot 3B085) due to being non-sterile. The affected lot passed sterility testing, however, another lot within the same cleaning cycle (Lot 3B087) failed sterility testing, exposing the affected lot to potential contamination. The use of non-sterile saline may cause serious adverse health consequences including difficulty breathing, general discomfort, nausea, wheezing, and infection. Infection can lead to life-threatening sepsis, and death. There have been no reports of injuries or death associated with this issue."
FDA MD Recalls | Proclaim XR 5 IPG Model 3660 (previously known as the ‘Proclaim 5 Elite IPG Model 3660’) Proclaim XR 7 IPG Model 3662 (Previously known as the ‘Proclaim 7 Elite IPG Model 3662’) Proclaim Plus 5 IPG Model 3670 Proclaim Plus 7 IPG Model 3672 Proclaim DRG IPG Model 3664 Infinity 5 IPG Model 6660 Infinity 7 IPG Model 6662 | "The Abbott Proclaim neurostimulation systems delivers low-intensity electrical impulses to nerve structures. The Proclaim XR and Proclaim Plus implantable pulse generators (IPGs) are used for Spinal Cord Stimulation (SCS). The Proclaim DRG IPG is used for Dorsal Root Ganglion (DRG) stimulation. The Infinity IPGs are used for Deep Brain Stimulation (DBS). The device has the option to disable delivery of therapy when a patient receives magnetic resonance imaging (MRI). The Patient Controller (PC) device, an iPhone or iPod installed with the PC application (app), connects to the patient’s IPG wirelessly via Bluetooth. The PC device is used to place a patient’s IPG in MRI mode, which disables the delivery of therapy. When the MRI procedure is complete, the PC device is then used to deactivate the IPG’s MRI mode and allow therapy to resume." | "Abbott is recalling its Proclaim and Infinity IPGs due to complaints from patients who are unable to exit MRI mode. The Patient Controller (iPhone/iPod) may lose the ability to connect or communicate with its IPG while in MRI mode. Example situations where this has occurred include when the PC device’s iOS operating system was updated, the PC app was updated or deleted, and when the IPG was deleted from the list of available Bluetooth devices on the PC device. A Clinician Programmer is required to be paired to the IPG for initial programming. When available, a Clinician Programmer previously paired with the patient’s IPG can be used to exit MRI mode. If there is no previously paired Clinician Programmer available, or if the Clinician Programmer lost its Bluetooth connection to the IPG, then there is no alternative option to exit MRI mode. The use of the affected IPGs may require surgery to remove the device and replace it with a new device. There have been 186 reported incidents and 73 reported injuries. There have been no reports of death."
FDA MD Recalls | One-Way Valve, 22F x 22M | "The Mallinckrodt Manufacturing, LLC One-Way Valve, 22F x 22M is part of the INOmax Delivery System, which delivers nitric oxide (INOmax) therapy gas into the tubes between the ventilator and patient airway (breathing circuit). The One-Way Valve is used to prevent backward flow of nitric oxide gas into the INOmax DSIR injector module. The omission of One-Way Valve can also result in high nitric oxide delivery to the patient. The intended prevention of the backward flow and higher than prescribed drug delivery applies to both high frequency ventilators and spontaneous mask setup." | "Mallinckrodt is recalling its One-Way Valve, 22F x 22M due to devices not opening properly, which prevents or reduces the flow of ventilated air or oxygen. The use of the One-Way Valve, 22F x 22M may cause serious adverse health consequences, including interruption in ventilation which may delay delivery of nitric oxide gas. This delay could lead to bradycardia in neonatal or pediatric patients. For all patients, interruption in ventilation may lead to insufficient oxygen in the blood resulting in respiratory failure, and death. There have been two complaints, and no reports of injuries or death."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid and Rescue Intra-Aortic Balloon Pump (IABP) devices are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave IABPs are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Getinge/Maquet/Datascope is recalling their Cardiosave Hybrid and Rescue Intra-aortic Balloon Pumps (IABPs) because they may overheat when the device’s internal temperature exceeds 176 degrees Fahrenheit. When this happens, the device will generate a “System Over Temperature” alarm that puts the device in standby mode. The pump then stops providing therapy. Interruptions in device therapy due to an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Between January 1, 2021 and June 2, 2023, Getinge/Maquet/Datascope has reported 213 complaints, including four injuries and one death."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid and Rescue Intra-Aortic Balloon Pump (IABP) devices are electromechanical systems used to inflate and deflate intra-aortic balloons (IAB). These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave IABPs are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Getinge/Maquet/Datascope is recalling their Cardiosave Hybrid and Rescue IABPs because an Autofill Failure alarm may put the device in standby mode resulting in therapy interruption. When IABP therapy begins, the IABP automatically fills the balloon catheter circuit with helium. Every two hours, or during changes to atmospheric pressure, the IABP autofill process occurs. If any step of the autofill process fails or does not complete, the device will generate an “Autofill failure” alarm message and the IABP pump will stop providing therapy. If the alarm persists, the patient may experience a prolonged interruption of therapy, which will require the healthcare provider to obtain a different console to continue therapy. If another IABP console is not available for use, alternative means of providing hemodynamic support (vasopressors, inotropes or alternate therapies) may be initiated by a healthcare provider as a temporizing measure. Interruption in device therapy due to an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Between January 1, 2021 and June 2, 2023, Getinge/Maquet/Datascope has reported 298 complaints related to this device issue, including two deaths and up to three injuries."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid and Rescue Intra-Aortic Balloon Pump (IABP) devices are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Getinge/Maquet/Datascope is recalling their Cardiosave Hybrid and Rescue IABPs because the pumps may stop unexpectedly due to Gas Gain or Gas Loss alarms. The alarm will interrupt therapy and require user intervention to correct and restart. If the alarm persists, the device must be switched to a back-up IABP to continue therapy. If no back-up IABP is available, heath care professionals must seek an alternative way of providing hemodynamic support. This can include temporary use of vasopressors, inotropes, or other alternative therapies. The alarm can be caused by patient movement (such as coughing, general movement, and swallowing); blood in the circuit; or loose connections, kinks, or abrasions to the tubing/catheter. Additionally, patients with fever or rapid heartbeat may experience a higher rate of gas loss, which can cause the alarm. Interruptions in device therapy due to an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Between January 1, 2021 and June 2, 2023, Getinge/Maquet/Datascope has reported 198 complaints related to this device issue, including three injuries and three devices involved in two patient deaths."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid and Rescue Intra-Aortic Balloon Pump (IABP) devices are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave IABPs are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Getinge/Maquet/Datascope is recalling their Cardiosave Hybrid and Rescue IABPs due to unexpected shut downs related to incorrect docking of the IABP console into the cart (Hybrid Mode). If the Cardiosave IABP is not docked into the cart correctly, it will not receive AC power and will unexpectedly run on battery power (Rescue Mode) and will shut down upon battery depletion. Additionally, if the Cardiosave IABP is not docked into the cart correctly (Hybrid Mode), the Cardiosave will not be able to charge the battery(ies) inserted and may affect the ability to safely transport a patient. Interruptions in device therapy due to an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Between January 1, 2021 and June 2, 2023, Getinge/Maquet/Datascope has reported 533 complaints, including one death."
FDA MD Recalls | HAMILTON-C1 with SW version <2.2.10 and neonatal option installed HAMILTON-C2 with SW version <2.2.5 and neonatal option installed HAMILTON-C3 with SW version <2.0.9 and neonatal option installed HAMILTON-T1 with SW version <2.2.10 and neonatal option installed | "Hamilton ventilators are used for people who require full or partial breathing assistance from a mechanical ventilator. Health care professionals use them during patient transport within or outside hospital settings. The ventilators support adults, children, and optionally infants and newborns." | "Hamilton Medical Inc. is recalling the HAMILTON-C1, C2, C3, T1 ventilators after receiving reports of software issues that may cause the ventilators to stop without notice. The ventilator will switch to 'Ambient State' if the device is used for more than 91 days without a restart in neonatal patient group. This may require a health care professional to intervene, provide air supply to the patient with a manual resuscitating device, and use an alternative device. Serious consequences or death may result for patients who do not breathe on their own or do not receive immediate intervention. There have been 80 reported complaints, no injuries, and no deaths related to this recall."
FDA MD Recalls | Carina Sub-Acute Care Ventilator | "The Carina Sub-Acute Care Ventilator is used for people who require full or partial breathing assistance from a mechanical ventilator inside a hospital. The ventilator can be used for invasive (trachea tube) or non-invasive (mask) ventilation." | "Draeger Medical, Inc. is recalling its Carina VentilatorsExternal Link Disclaimer due to the presence of contaminants in the device’s airpath. The contaminants exceed acceptable levels if used by pediatric patients for more than 30 days. During testing, the firm found polyether polyurethane (PE-PUR) in the airpath of the Carina ventilator. This results in the emission of contaminant 1,3-Dichloropropan-2-ol, a potential carcinogen. The use of Carina Sub-Acute Care Ventilators in pediatric patients may cause serious adverse health consequences, including exposure to hazardous chemicals, toxic reactions, and death. There have been no reported deaths, complaints, or incidents associated with this issue."
FDA MD Recalls | Impella RP Flex with SmartAssist | "The Impella RP Flex with Smart Assist System Catheter is used for up to 14 days in patients that develop acute right heart failure after left ventricular assist device implantation. The device is placed via the internal jugular vein and supports the right chamber of the heart (ventricle) by pumping blood into the pulmonary artery." | "Abiomed is recalling the labeling for Impella RP Flex with Smart Assist System Catheter because the catheters' Instructions for Use (IFU) do not appropriately address precautions for health care providers to take when treating patients whose anticoagulation clotting time is below the recommended value. Patients with central venous lines and cardiac cannulas with systemic anticoagulation below IFU recommendation of 160-180 seconds are most at risk. Please note that clinicians may continue to use the devices. The use of affected catheters may cause serious adverse health consequences including the risk of blood clots or particle deposits forming or death. There have been 12 reported injuries. There have been no reports of death."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and the Cardiosave Rescue IABP are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Datascope/Maquet/Getinge is recalling the Cardiosave Hybrid and Rescue Intra-aortic Balloon Pumps (IABPs) because they may shutdown unexpectedly due to failures in the printed circuit board assembly (PCBA) in the charging path. If this occurs, the device may lose the ability to charge batteries in one or both battery bays/slots due to electrical surge damage in the Power Management Board Charging Path Circuit. Using an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Datascope/Maquet/Getinge reported 252 complaints. There have been no reports of injuries or death."
FDA MD Recalls | Philips Respironics Trilogy Evo, Trilogy Evo O2, Trilogy EV300, Trilogy Evo Universal | "Philips Respironics Trilogy Evo, Evo O2, EV300, and Evo Universal ventilators provide continuous positive airway pressure (CPAP) and intermittent positive airway pressure breathing support for people who need mechanical ventilation. They are intended for pediatric through adult patients weighing at least 2.5 kg and are used in hospitals and health care settings and for patient nonemergency transport. The Trilogy EV300 and Trilogy Evo may also be used by people requiring mechanical ventilation in their homes." | "Philips is recalling Trilogy Evo, Evo O2, EV300, and Evo Universal ventilators after detecting dust and dirt from the environment in the air path of some devices. Extended exposure to environmental contaminants such as dust and dirt can lead to buildup that may block air vents and cause the device to stop delivering the right amount of air pressure, or air volume/flow. If the ventilator fails to provide the right level of breathing support, patients may not receive enough oxygen (hypoventilation) and may experience a build-up of carbon dioxide or other gas pressure, which can lead to serious injury or death. Philips has received 542 reports about this issue. There are currently two reported injuries and one death."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and the Cardiosave Rescue IABP are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Datascope/Maquet/Getinge is recalling the Cardiosave Hybrid and Rescue Intra-aortic Balloon Pumps (IABPs) because they may shutdown unexpectedly due to electrical failures in the Power Management Board or Solenoid Board (Power Source Path). Using an affected pump may cause serious adverse health events, including unstable blood pressure, injury (for example: inadequate blood supply or a vital organ injury), and death. Datascope/Maquet/Getinge reported 26 complaints. There have been no reports of injuries, or death."
FDA MD Recalls | SIGMA Spectrum Infusion Pump with Master Drug Library (Version 8) and Spectrum IQ Infusion System with Dose IQ Safety Software (Version 9) | "The SIGMA Spectrum Infusion Pump with Master Drug Library (Version 8) and Spectrum IQ Infusion System with Dose IQ Safety Software (Version 9) are software-controlled infusion pumps that deliver controlled amounts of fluids such as pharmaceutical drugs, blood and blood products, and other required therapies. The fluids are provided through infusion into a vein or other cleared route. These infusion pumps are used in hospitals and other healthcare facilities." | "Baxter is recalling SIGMA Spectrum Infusion Pumps with Master Drug Library (Version 8) and Spectrum IQ Infusion Systems with Dose IQ Safety Software (Version 9) due to increased reports of false alarms for upstream occlusion after pump software upgrades to version v8.01.01 and v9.02.01, respectively. False upstream occlusion alarms cause interrupted or delayed therapy and contribute to clinician fatigue, which may cause serious adverse health consequences, especially for people receiving life-sustaining medications. Use of these products may lead to serious injury or death. Baxter reports 131 complaints, three serious injuries, and no deaths related to this issue."
FDA MD Recalls | TruSignal Adult Pediatric Sensor TruSignal AllFit Sensor TruSignal Sensitive Skin Sensor TruSignal Wrap Sensor TruSignal Ear Sensor TruSignal Integrated Ear Sensor with GE Connector TruSignal Integrated Ear Sensor with Datex Connector TruSignal Integrated Ear Sensor with Datex Connector TruSignal Integrated Ear Sensor with Ohmeda Connector | "TruSignal sensors continuously monitor the amount of oxygen found in the blood flowing through the arteries (arterial oxygen saturation or SpO2) and pulse rate through a sensor placed on the skin. The sensors are designed and tested to help ensure patients have enough oxygen in their blood to keep their bodies functioning and to give care providers information to make treatment decisions." | "GE HealthCare is recalling malfunctioning TruSignal sensors that may reduce the amount of energy sent to the heart during defibrillation without any notification to the care provider, which could prevent delivery of lifesaving therapy in a critical situation. This issue is most hazardous to hospitalized patients who may need defibrillation for cardiac arrest. Affected sensors may also unintentionally expose patients to electrical currents from other sources or may provide inaccurate measurements of SpO2, which can impact treatment decisions. Use of defective TruSignal SpO2 Sensors may cause serious injuries or death. GE HealthCare has not received any reports of injuries or death for this issue."
FDA MD Recalls | Impella 5.0 Blood Pump, Product Number 005062 Impella CP Blood Pump, Product Number 0048-0032 Impella 2.5 Blood Pump, Product Number 005042 Impella CP with SmartAssist Blood Pump, Product Numbers 0048-0024, 0048-0045, 1000080 Impella LD Blood Pump, Product Number 005082 Impella 5.5 with SmartAssist Blood Pump, Product Numbers 0550-0008 And 1000100 | "Impella Left Sided Blood Pumps are used for short term support of the pumping chambers of the heart (ventricles) during high-risk catheter-based procedures called percutaneous coronary interventions (PCI). Impella Left Sided Blood Pumps also are used when there is ongoing cardiogenic shock that occurs less than 48 hours after a severe heart attack (acute myocardial infarction), open-heart surgery, or when the heart is not functioning well due to a condition called cardiomyopathy. Impella therapy aims to reduce the work of the heart’s ventricles and provide support for the circulatory system so the heart has time to recover. " | "Abiomed is recalling the Impella Left Sided Blood Pumps because the pumps’ Instructions for Use (IFU) do not adequately address precautions to take when treating patients who have undergone transcatheter aortic valve replacement (TAVR). The IFU lacks guidance to clinicians on how to manage use of Impella in patients with TAVR and fails to describe how the issue may present if an Impella interacts with TAVR. There is a potential risk that the Impella motor housing may come into contact with the distal stent of a transcatheter aortic valve replacement (TAVR). The contact may damage or destroy the motor’s impeller blades. The damaged Impella system may have reduced blood flow or pump stop, which may delay therapy or fail to provide enough support to the patient. This could be life threatening in people who require high levels of support. There is also a risk that pieces of the broken blades could enter the patient’s bloodstream. Use of affected devices may cause serious injury or death. Abiomed reports 30 complaints, 26 injuries, and four deaths related to this issue. Please note that clinicians may continue to use the devices, with the additional instructions for patients with TAVR in mind."
FDA MD Recalls | Amplatzer Steerable Delivery Sheath | "The Abbott Amplatzer Steerable Delivery Sheath is a cardiac catheter inserted through the skin that provides a pathway for catheter-based devices to be introduced into the chambers of the heart. It is specifically used to deliver the Amplatzer Amulet Left Atrial Appendage Occluder." | "Abbott is recalling the Amplatzer Steerable Delivery Sheath due to an increased risk for air bubbles (air emboli) to be introduced into patients who have procedures with this device. Air embolism can lead to injuries such as sudden (acute) reduction in blood flow to the heart (indicated by ST elevation), fast or slow heartbeat (tachycardia or bradycardia), low blood pressure (hypotension), and lack of enough oxygen in the blood (desaturation). Air embolism can also potentially lead to serious health consequences including stroke and death. Abbott reports 26 incidents, 16 injuries, and no deaths related to this issue."
FDA MD Recalls | Cobalt XT, Cobalt, Crome ICDs and CRT-Ds Claria MRI, Amplia MRI, Compia MRI, Viva, Brava CRT-Ds Visia AF, Visia AF MRI, Evera, Evera MRI, Primo MRI, Mirro MRI ICDs | "Medtronic’s implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) are implantable cardiac devices that are intended to monitor and regulate heart rate and rhythm. These devices automatically detect and treat life-threatening irregular heartbeats (arrythmia) by producing an electric shock (cardioversion or defibrillation) to restore a normal heartbeat. Some devices can also provide cardiac resynchronization therapyExternal Link Disclaimer to treat heart failure." | "Medtronic is recalling all ICDs and CRT-Ds, manufactured after 2017, with a glassed feedthrough, as they may deliver low or no energy output when high voltage therapy is needed due to inappropriate activation of the Short Circuit Protection (SCP) feature. The issue is more likely to occur for devices with a glassed feedthrough that are configured to deliver therapy in the AX>B delivered pathway. A reduced-energy shock, or no shock at all, may fail to correct a life-threatening arrhythmia, which can lead to cardiac arrest, other serious injury, or death. There are additional risks of harm if a patient with one of these devices needs additional surgical procedures to remove and replace the device. Medtronic reports 28 incidents, 22 injuries, and no deaths for this issue."
FDA MD Recalls | Quidel Triage Cardiac Panel Quidel Triage Cardiac Panel, Troponin I Quidel Triage Cardiac Panel, CK-MB and Troponin I | "The Quidel Triage Cardiac Panel is a blood test that detects amounts of specific enzymes (creatine kinase MB or CK-MB) and proteins (myoglobin and troponin I) in blood or plasma samples. The test is used to help doctors quickly diagnose when someone is experiencing a heart attack (myocardial infarction) and other life-threatening injuries to the heart muscle." | "Quidel Cardiovascular Inc. (QuidelOrtho) is recalling the Quidel Triage Cardiac Panel after reports of inaccurate tests showing lower than expected troponin levels in samples. A falsely low or false negative troponin level may cause a delayed or missed diagnosis of myocardial infarction — a serious, potentially fatal condition that requires the earliest detection possible to treat and protect the heart muscle effectively. This poses a particular risk to people who experience a silent myocardial infarction with no other symptoms, those who have unusual or atypical symptoms, and those with non-ST-segment elevation myocardial infarction (NSTEMI). Continued use of these products without appropriate mitigations could cause serious injury or death. QuidelOrtho reports 41 complaints, no injuries, and no deaths related to this issue."
FDA MD Recalls | NOxBOXi Nitric Oxide Delivery System | "The NOxBOXi Nitric Oxide Delivery System (NOxBOXi) delivers a constant concentration of nitric oxide (NO) medical gas for inhaled nitric oxide therapy through the ventilator that provides breathing support to the patient. The system can also monitor nitric oxide, nitrogen dioxide, and oxygen levels while the patient receives NO therapy. It is approved for use in newborns, or neonates, including those with congenital heart disease." | "NOxBOX Ltd. is recalling NOxBOXi Nitric Oxide Delivery Systems due to a misalignment of the check valve in the device’s manifold, which may cause the manifold to fail. If it fails, nitric oxide or oxygen may leak out. The device may also change gas cylinders earlier than expected or not at all. This issue may lead to interrupted nitric oxide therapy, causing a drop in blood oxygen levels (desaturation) or increased pressure in the artery that carries oxygen from the heart to the lungs (pulmonary artery) which may pose even greater risks to neonates with congenital heart disease. Use of affected devices could lead to serious injury or death in an already fragile patient population. NOxBOX Ltd. reports no injuries and no deaths related to this issue."
FDA MD Recalls | Oxylog 3000 Plus Emergency and Transport Ventilator | "The Oxylog 3000 Plus Emergency and Transport Ventilator is used for people who require full or partial breathing assistance from a mechanical ventilator. Healthcare professionals use it during patient transport, either in an ambulance or aircraft or to move people using ventilators throughout the hospital and recovery room." | "Draeger Medical is recalling the Oxylog 3000 Plus Emergency and Transport Ventilator after receiving reports that the device stopped ventilation because of a depleted battery, even after being re-connected to AC power. The ventilator may not automatically switch back to using AC power when it is plugged in and may continue using the battery until it is depleted, then stop providing ventilation. A battery alarm: “Charge int. battery” and “Int. battery discharged” does occur with this issue. Stopped ventilation may cause difficulty breathing (respiratory distress), lack of oxygen (hypoxia), slow heartbeat (bradycardia), a sudden stop of the heart (cardiopulmonary arrest), other severe injuries, or death. Draeger Medical reports six complaints, no injuries, and no deaths related to this issue."
FDA MD Recalls | MEGADYNE MEGA 2000 Patient Return Electrode 0800 MEGADYNE MEGA SOFT Reusable Patient Return Electrode 0830 MEGADYNE MEGA SOFT Dual Reusable Patient Return Electrode 0835 MEGADYNE MEGA SOFT Pediatric Patient Return Electrode 0840 MEGADYNE MEGA SOFT Universal Patient Return Electrode 0845 MEGADYNE MEGA SOFT Universal Dual Patient Return Electrode 0846 MEGADYNE MEGA SOFT Universal Plus Patient Return Electrode 0847 MEGADYNE MEGA SOFT Universal Plus Dual Patient Return Electrode 0848 | "MEGA 2000 (Figure 1) and MEGA SOFT (Figure 2) Reusable Patient Return Electrodes are soft pads used during electrosurgery. In electrosurgery, an electric current is used to heat or cut tissue or to stop bleeding. The electric current is generated by an electrosurgical generator and delivered to the tissue with a small pen-like attachment. A return electrode pad contacts the skin of the laying patient during use and conducts the electric current from the patient's tissue back to the electrosurgical unit, or generator, to reduce the risk of excessive heating." | "Megadyne is recalling MEGA 2000 and MEGA SOFT Reusable Patient Return Electrodes after receiving reports of patients receiving burn injuries from use in both pediatric and adult patients. These burns may be as serious as third-degree burns requiring medical intervention and may lead to a longer hospital stay, scarring, and potentially more surgeries for pediatric and adult patients. Megadyne is continuing to evaluate the root cause for patient burns. Use of affected products may cause serious injury or death. Megadyne reports 63 injuries and no deaths related to this issue. If additional actions are necessary to further help ensure the safe use of these devices, Megadyne will notify customers."
FDA MD Recalls | ARROW Endurance Extended Dwell Peripheral Catheter System | "The ARROW Endurance Extended Dwell Peripheral Catheter System allows health care providers access to a patient’s peripheral vascular system, for short term use, to sample blood, monitor blood pressure, or administer fluids, blood, and blood products." | "Teleflex, and their subsidiary Arrow International, are recalling the ARROW Endurance Extended Dwell Peripheral Catheter System after reports of catheter separation or leakage. If the catheter separates while in a blood vessel, the catheter fragments could be left in the bloodstream and may migrate to other places in the body. This issue may cause serious injury, including blockage of blood vessels, inadequate blood flow, injury to blood vessel walls, blood clots, blockage of the lung arteries (pulmonary embolism), heart attack, or death. Teleflex/Arrow International reports 83 complaints related to this issue. There have been 18 reported injuries and no deaths."
FDA MD Recalls | Walnut Wearable Smart Thermometer | "The rechargeable Walnut Wearable Smart Thermometer is a reusable thermometer that is paired with an application on a mobile phone. It is intended for home use and continuously monitors the body temperature of infants and children 0 to 6-years-old." | "BearCare, Inc. is recalling the rechargeable Walnut Wearable Smart Thermometer after receiving reports of skin burns and irritation resulting from use of the device. Potential overheating and leakage of corrosive chemicals from the device could cause skin irritation, burns, other serious injuries, or death. BearCare, Inc. reports five complaints related to this issue."
FDA MD Recalls | Impella 5.5 with SmartAssist | "The Impella 5.5 with SmartAssist System is used for up to 14 days to support the pumping chambers of the heart (ventricles) when there is ongoing cardiogenic shock that occurs less than 48 hours after a severe heart attack (acute myocardial infarction), open-heart surgery, or when the heart is not functioning well due to a condition called cardiomyopathy." | "Abiomed is recalling specific Impella 5.5 with SmartAssist sets after receiving customer complaints about purge fluid leaking from the purge sidearm of the pump. If a purge leak occurs the system will experience low purge pressures, prompting alarms and requiring evaluation. If the leak issue is not resolved, persistent low purge pressure and purge flow may lead to pump stop and loss of therapy. In patients who are critical, failure of the pump’s support can lead to further deterioration and worsening of their already critical condition and may even lead to serious injury or death. Abiomed reports 179 complaints, three injuries and no deaths related to this recall."
FDA MD Recalls | Pilot COVID-19 At-Home Test | "The SD Biosensor Pilot COVID-19 At-Home Test is a lateral flow immunoassay device intended for the qualitative detection of the nucleocapsid protein antigen found in the SARS-CoV-2 virus. This test is authorized for non-prescription home use, self-collected anterior nasal (nares) swab samples from individuals aged 14 years or older, or adult collected anterior nasal (nares) swab samples from individuals aged 2 years or older. " | "SD BioSensor, Inc. is recalling certain Pilot COVID-19 At-Home Test kits, distributed by Roche Diagnostics in U.S. Roche packaging, after finding that the liquid solution in the test kit may be contaminated with bacteria such as Enterococcus, Enterobacter, Klebsiella and Serratia. Infection from these bacteria species may cause illness in people with weakened immune systems or those with direct exposure to the contaminated liquid solution through standard handling, accidental spills, or misuse of the product. In addition to the risk of infection, this contamination may impact the performance of the test, and false results may occur. Use of the affected product may cause serious adverse health consequences or death. SD Biosensor, Inc. has received no complaints or reports of injuries, deaths, or adverse events."
FDA MD Recalls | Seattle PAP Plus, VentStar Basic (P)180, VentStar Basic(P)250, VentStar Watertrap (P) 180, VentStar Watertrap (P) 180 w/oLL, VentStar Anesthesia WT (P)180, VentStar breathing bag Set (P) 110, ID Circuit Basic (P) 180, VentStar Anesthesia (N) 180, VentStar Basic (N) 180, VentStar Watertrap (N), 180 VentStar Bag Set (N) 110, VentStar Coax (P) 150, Anesthesia Circuit Kit Coax 1, Anesthesia Circuit Kit Coax 3, Ventstar Coax, Ventstar Coax 180, ID Coax 180, Ventstar Coax 230, Anesthesia Set Coax 180, Set2Go Ventilation 12 (A), and Anesthesia Circuit Kit Coax HEPA | "The Draeger Medical Seattle-Positive Airway Pressure (PAP) Plus system is intended to provide Bubble continuous positive airway pressure (CPAP) therapy that supports infants who are struggling to breathe (in respiratory distress) while in the hospital. The other products included in this recall are breathing circuit and/or anesthesia kits used together with ventilators during surgery or in the intensive care unit to support the breathing of infants, children, and adults." | "Draeger Medical is recalling the Seattle PAP Plus as well as VentStar and other breathing circuit/anesthesia kits after finding that glued connections may loosen before or during ventilation due to a manufacturing error. Partial or complete detachment of components including the water trap, y-piece, or hose connector, may occur as a result. Loosening or detachment of these parts can interrupt the breathing circuit and may cause severe injury including lack of oxygen (hypoxia) or death. The risk of injury or death is especially great for critically ill patients including newborns (neonates). Draeger Medical reports no injuries or deaths related to this issue."
FDA MD Recalls | Replacement Batteries for Plum 360, Plum A+, and Plum A+3 Infusion Systems | "The Plum 360, Plum A+ and Plum A+3 are large volume infusion pumps used to give fluids to patients in precisely controlled infusion amounts and rates. The system delivers blood or blood products, drugs, and other fluid mixtures through subcutaneous, intramuscular, intravenous, and intrathecal administration. The batteries and replacement batteries are used when the pump is not plugged into AC power, for example when a patient is being transported." | "ICU Medical is recalling replacement batteries for the Plum 360, Plum A+, and Plum A+3 Infusion Systems because a manufacturing defect has substantially diminished how long they can be used to run the system. If the pump is running on battery power and there is no AC power backup available, the system may shut down an ongoing infusion and power down sooner than expected. This issue may cause serious injury or death to patients due to interruption, under-infusion, or delays in the delivery of critical fluids, blood products, and medications. ICU Medical has received 54 (Plum 360) and 465 (Plum A+/A+3) complaints regarding this issue and has not reported any injuries or deaths. However, the FDA has identified an adverse event that may be related to the recall issue."
FDA MD Recalls | Ivenix Infusion System | "The Ivenix Infusion System is a large volume pump used in hospitals and other health care settings to give fluids to patients in precisely controlled infusion amounts and rates. The system delivers blood or blood products, drugs, and other fluid mixtures through blood vessels, the back/spine (epidural), or beneath the skin (subcutaneous). It is used for adults and pediatric patients I.e.neonates, infants, children and adolescents." | "Fresenius Kabi USA, LLC is recalling the Ivenix Infusion System due to a leak in the system that allows fluid to enter the Administration Set Loading Area near the Air Detector. The leak damages the electrical system, leading to loss of power and failure of the system’s Set Identification (Set ID) sensor. Depending on the timing of the failure, this can delay or interrupt treatments. This issue may cause serious injury or death to patients due to interruption, under-infusion, or delays in the delivery of critical fluids, blood products, and medications. Fresenius Kabi USA has received 14 complaints related to this issue. There have been no reported injuries or deaths."
FDA MD Recalls | Philips Respironics DreamStation1 (Uno Remediated Devices) | "Philips Respironics (Philips) DreamStation devices help people with breathing conditions keep breathing at a regular rhythm. They are used both in hospitals and health care settings as well as by people in their homes. The DreamStation can provide both continuous positive airway pressure (CPAP) and bilevel positive airway pressure (BiPAP) support. A CPAP machine keeps your airway open by providing a stream of air at a continuous pressure through a mask. CPAP machines are often prescribed to people with obstructive sleep apnea to keep their airways open during sleep. A BiPAP machine pumps air under varying pressure into the airway of the lungs. BiPAP machines use a higher pressure when you breathe in and lower pressure when you breathe out. These devices are also prescribed to people with obstructive sleep apnea to keep their airways open during sleep." | "Philips is recalling certain reworked Philips DreamStations because some devices were assigned incorrect or duplicate serial numbers during initial programming. This duplication can cause therapy to be delivered using the wrong prescription or factory default settings. Additionally, it may fail to deliver any therapy at all. There is no warning or indication to the user that the DreamStation is not working the way the doctor intended or prescribed. Incorrect therapy or therapy failure may lead several health conditions such as respiratory failure, heart failure, serious injury, and death. Philips has received 43 complaints about this issue. There are currently no reported injuries or deaths."
FDA MD Recalls | FreeStyle Libre Flash Glucose Monitoring System, FreeStyle Libre 14 day Flash Glucose Monitoring System, FreeStyle Libre 2 Flash Glucose Monitoring System | "The FreeStyle Libre, Libre 14 day, and Libre 2 Flash Glucose Monitoring Systems are intended to provide continuous monitoring of glucose levels. These devices help people manage diabetes by detecting trends and tracking patterns in glucose levels so treatment can be adjusted as needed. They are indicated for single patient use, and require a prescription." | "Abbott is recalling the FreeStyle Libre, Libre 14 day, and Libre 2 Flash Glucose Management Systems’ reader devices, which use rechargeable lithium-ion batteries, may get extremely hot, spark, or catch on fire if not properly stored, charged, or used with its Abbott provided USB cable and power adapter. This does not affect any of the FreeStyle Libre family of sensors. The potential for overheating, spark, or fire may occur when charging the Reader with non- Abbott adapters or non-Abbott USB cables combined with misuse of the Reader and its components. Examples of misuse include exposure to liquids, damage, and introduction of foreign material into the ports. The Abbott-provided USB cable and power adapter limit the power provided to safely charge the battery, whereas USB cables and power adapters manufactured by a third party may allow much higher power, increasing the risk of overheating, spark, or fire. The Reader, if not properly stored, charged, or used with its Abbott provided USB cable and power adapter, may expose users to extreme heat and/or fire which can cause serious injuries or death. Additionally, users may delay or miss a critical diabetes treatment if the system cannot be used after is damaged by extreme heat or fire. There have been 88 incidents, including at least seven reports of fires, one injury, and no deaths involving this issue. "
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and the Cardiosave Rescue IABP are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Datascope, a subsidiary of Getinge, is recalling Cardiosave Hybrid IABPs and Rescue IABPs because a communication loss between the executive processor printed circuit board assembly (PCBA) and the video generator PCBA may cause an unexpected shutdown. This shutdown will occur without warnings or alarms to alert the user. Once the error occurs, one or both of the PCBAs will need to be replaced for the device to work again. An unexpected pump shutdown and any interruption to therapy that occurs can lead to unstable blood flow (hemodynamic instability), organ damage and/or death, especially for people who are critically ill and most likely to receive therapy using these devices. Datascope/Getinge reported 42 complaints about PCBA communication failures resulting in unexpected shutdowns. There are no reports of any injuries or deaths related to this issue."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and the Cardiosave Rescue IABP are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Datascope, a subsidiary of Getinge, is recalling Cardiosave Hybrid IABPs and Rescue IABPs because the coiled cable connecting the display and base on some units may fail, causing an unexpected shutdown. This shutdown will occur without warnings or alarms to alert the user. Once the error occurs, the coiled cord needs to be replaced for the device to work again. The issue may occur in units distributed before July 24, 2017. An unexpected pump shutdown and any interruption to therapy that occurs can lead to unstable blood flow (hemodynamic instability), organ damage and/or death, especially for people who are critically ill and most likely to receive therapy using these devices. From June 2019 to August 2022, Datascope/Getinge has reported 44 complaints about damaged coiled cords resulting in unexpected shutdowns. There have been no reports of injuries or deaths related to this issue."
FDA MD Recalls | BALLARD ACCESS Closed Suction System for Neonates/Pediatrics with Y-Manifold; BALLARD ACCESS Closed Suction System for Neonates/Pediatrics with Elbow Manifold | "The manifold of the BALLARD ACCESS Closed Suction System for Neonates/Pediatrics is used to provide access to the artificial airway of a child, infant, or neonate, without breaking the ventilation circuit. It is a single patient use device that is used by trained medical professionals. The BALLARD ACCESS Closed Suction manifold is indicated for use together with the BALLARD ACCESS Closed Suction catheter to remove secretions from the artificial airway of children, infants, and neonates" | "Avanos Medical Inc. is recalling the Avanos BALLARD ACCESS Closed Suction System for Neonates/Pediatrics with Y-Manifold or Elbow Manifold (labeled 72-hour-use) after receiving complaints from customers about cracked manifolds during use. A cracked manifold can cause leaks in the respiratory circuit, which may lead to the patient receiving inadequate ventilation. Inadequate ventilation and oxygenation can lead to lead serious brain damage or death, especially in the vulnerable patient population (children, infants, and neonates) who receive care using this product. Additional risks from this issue include the possibility that a foreign body could be introduced through the cracked manifold, causing infection, damage to the airway, or a blocked endotracheal tube. Use of this product could lead to serious injury or death. Avanos Medical, Inc. reports four injuries associated with this issue."
FDA MD Recalls | Philips Respironics Trilogy 100, Trilogy 200, and Garbin Plus ventilators | "Philips Respironics (Philips) Trilogy and Garbin ventilators are used to provide breathing assistance to both pediatric and adult patients. A ventilator is intended to mechanically control or assist patient breathing by delivering a predetermined percentage of oxygen. Trilogy and Garbin ventilators are intended to be used in homes and health care settings, including when patients may be using a wheelchair or gurney." | "Philips is recalling certain reworked Philips Trilogy 100, Trilogy 200, and Garbin Plus ventilators for two recent issues: The silicone sound abatement foam installed to replace PE-PUR foam may separate from the plastic backing due to adhesive failure. The silicone foam material may potentially move and block the airpath, which may reduce air flow in the ventilator and could also cause the device to alarm. If an alarm is not recognized or acted upon, the patient could experience trouble breathing, such as asphyxia, hypoventilation, or hypoxemia, which can be life threatening. Philips observed residual PE‐PUR sound abatement foam in some reworked ventilators that were returned to customers. Further exposure to PE-PUR foam may cause potential health risks, which can result in serious injury, cause permanent impairment, and require medical intervention to prevent permanent injury to users. As of January 4, 2023, there have been 82 MDRs (63 MDRs for foam delamination and 19 MDRs for particulate debris contamination) and no deaths."
FDA MD Recalls | Nuclear Medicine 600/800 Series systems | "GE Nuclear Medicine 600 and 800 Series systems are used by health care professionals to evaluate diseases, trauma, abnormalities, and disorders. These systems create images of the body to help health care providers assess organ function or detect and diagnose issues such as cardiovascular disease, neurological disorders, and cancer." | "GE HealthCare is recalling the Nuclear Medicine 600 and 800 Series systems after identifying an issue with two mechanisms that prevent uncontrolled detector movement. First, there is a risk that the ball screw that serves as design mitigation for the suspended mass of the detector may fail. Additionally, some devices are also missing a safety key that should prevent the detector from a catastrophic fall when the ball screw fails. If the ball screw fails and the safety key is missing, the 1212-pound (550-kg) detector could fall, potentially crushing or trapping a patient, which may result in serious injury or death. There have been no complaints, no injuries, and no deaths reported related to this issue. This issue was identified internally by GE HealthCare. While conducting this recall, GE HealthCare has identified and corrected eight instances of the missing safety key."
FDA MD Recalls | Skippack Medical Lab SARS-CoV-2 Antigen Rapid Test (Colloidal Gold) | "The Skippack Medical Lab SARS-CoV-2 Antigen Rapid Test (Colloidal Gold) uses patient samples to detect proteins, called antigens, found on the SARS-CoV-2 virus. The samples are collected in one of two ways: Nasal (anterior nares) swab Nasopharyngeal swab, which reaches deep inside the nose, reaching the back of the throat or from the middle part of the throat (pharynx)" | "Universal Meditech Inc. is recalling Skippack Medical Lab COVID-19 Direct Antigen Rapid Tests (Colloidal Gold) because these tests were distributed to U.S. customers without authorization, clearance or approval from the FDA. Neither Universal Meditech Inc. nor SML Distribution LLC, has provided the FDA with adequate data to show that the test's performance is accurate. This means there is a risk of potential false negative, false positive, or misinterpretation of results. Use of the affected product may cause serious adverse health consequences or death. Universal Meditech Inc. has received no complaints or reports of injuries, deaths, or adverse events."
FDA MD Recalls | CADD infusion system Administration Sets and Cassette Reservoirs | "CADD ambulatory infusion systems deliver controlled amounts of medications to patients in hospital, outpatient, and home settings. The CADD Administration Sets and CADD Medication Cassette Reservoirs are disposable and intended for use with CADD infusion pumps. They deliver the medication into a patient’s vein or through other cleared routes of administration." | "Smiths Medical is recalling certain CADD Administration Sets and Medication Cassette Reservoirs for two potential issues: Tubing occlusion prevents delivery or causes under-delivery of medication. Manufacturing variations may cause the green CADD Flow Stop arm to compress and partially block (occlude) the tubing before clinical use. If the occlusion does not resolve when an affected reservoir or administration set is connected to the pump, the pump may not detect the occlusion, resulting in under-delivery or non-delivery of medication, though the pump will display that the infusion is running properly. False “No Disposable Attached (NDA)” alarm prevents pump use. Manufacturing variations on certain CADD Medication Cassette Reservoirs with Flow Stop may interfere with the pump detecting a properly attached CADD cassette. The pump will give an audible warning. If the audible warning is not resolved within two minutes, an NDA alarm will occur. The user must clear the alarm and resolve the cause of the NDA event before using the pump, resulting in the delay or interruption of therapy. Either of these recall issues can cause delay of therapy, interruption of therapy, or under-delivery of medication, which all have the potential to cause serious patient harm or death. Smiths Medical reports there have been 1,571 incidents, 14 injuries and 2 deaths related to the tubing occlusion issue and 9,101 incidents, 11 injuries and no deaths related to the false NDA alarm issue."
FDA MD Recalls | Mahurkar Acute Dual Lumen High Flow (13.5 French) Hemodialysis Catheters (Mahurkar QPlus) | "Mahurkar Acute Dual Lumen High Flow (13.5 French) Hemodialysis Catheters are implanted and used in hemodialysis (treatment to filter waste and water from the blood), apheresis (a technology that separates donated blood components to treat certain illnesses), and infusion (a method of putting fluids into the bloodstream) for up to 21 days." | "Medtronic is recalling the Mahurkar Acute Dual Lumen High Flow (13.5 French) Hemodialysis Catheters due to a potential catheter hub defect that may cause leaks across the catheter’s tubes. During treatment, this leak could result in mixing of the arterial and venous blood, lead to increased recirculation and poor dialysis, or cause the development of blood clots in the blood vessels. The use of the defective catheter may cause serious adverse health outcomes, including bleeding or the need for surgical removal and replacement of the affected catheter. There have been two injuries reported related to this issue and no reports of death."
FDA MD Recalls | LifeSPARC System | "The LifeSPARC System is intended to pump the blood through an extracorporeal circuit for periods lasting less than 6 hours for the purpose of providing either: Full or partial cardiopulmonary bypass (i.e., circuit includes an oxygenator) during open surgical procedures on the heart or great vessels; or Temporary circulatory bypass for diversion of flow around a planned disruption of the circulatory pathway necessary for open surgical procedures on the aorta or vena cava. The LifeSPARC System has two components: the LifeSPARC Pump and the LifeSPARC Controller. The LifeSPARC Pump is a single-use pump, and the LifeSPARC Controller provides the interface between pump and user, as well as the power and electrical signals to drive the pump." | "LivaNova (TandemLife) is recalling the LifeSPARC Controller, part of the LifeSPARC System due to a software update being available to address the previous software malfunction that may erroneously detect frozen or unresponsive software and trigger the device to enter Critical Failure mode—clearing the controller screen and issuing an alarm that cannot be muted or turned off. This recall is an addition to a previous recall for the same issue, LivaNova (TandemLife) Recalls LifeSPARC System for Risk of Unintentional Extended Pump Stop During Controller Critical Failure. The current recall provides information about a software update now available to address this issue. The FDA has received two injuries and no deaths due to this issue."
FDA MD Recalls | Emergent RSDL (Reactive Skin Decontamination Lotion) Kits | "The RSDL (Reactive Skin Decontamination Lotion) kit is an easy-to-open packet containing a sponge soaked with lotion. The RSDL kit is carried by members of the military and is intended to remove and/or neutralize chemical warfare agents and T-2 Toxin from the skin." | "Emergent is recalling RSDL (Reactive Skin Decontamination Lotion) kits from a specific batch after customers noted that some packets were leaking. Product in leaking packets may not perform as effectively as intended. The leak in the package could cause the lotion-bearing sponge to dry out or lose lotion content. This may reduce the effectiveness of the kit. Users also may be unnecessarily exposed to the lotion if the package is leaking. In a situation of exposure to chemical warfare agents or T-2 toxin requiring immediate decontamination with RSDL, the potential for reduced or absent device effectiveness could make the exposed individual vulnerable to health hazards from the chemical agent exposure. Additionally, unnecessary exposure to the non-opened packets that have been contaminated by the lotion from the leaked packets could result in unintentional prolonged skin, eyes, and mucous membrane exposure to the military personnel as well as unsuspecting handlers of the subject RSDL kits. Use of the affected kits may lead to serious injury and/or death. There have been 3 complaints and no reports of injuries or deaths related to this issue."
FDA MD Recalls | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | "The Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) and the Cardiosave Rescue IABP are electromechanical systems used to inflate and deflate intra-aortic balloons. These systems provide temporary support to the left ventricle through counter pulsation. Once the balloon is positioned in the aorta, the pump is set to work in synchrony with the electrocardiogram or arterial pressure waveform to make the balloon inflate and deflate at the right time during the cardiac cycle. Cardiosave Intra-Aortic Balloon Pumps are indicated for acute coronary syndrome, cardiac and non-cardiac surgery, or complications of heart failure in adults. They are used in health care facilities." | "Datascope, a subsidiary of Getinge, is recalling Cardiosave Hybrid IABPs and Rescue IABPs because a compromised intra-aortic balloon (a burst, leaking, or torn balloon) can cause blood to enter the IABP during therapy (a blood back event). Blood in the pump can cause the following issues: Unexpected pump shutdown if blood comes in contact with electrical components. A pump shutdown can lead to unstable blood flow (hemodynamic instability), organ damage and/or death, especially for people who are critically ill and most likely to receive therapy using these devices. If therapy continues, patients may have helium released into their blood. Helium gas bubbles (gas emboli) in the blood can damage organs (including the brain). Patient blood loss The user and/or subsequent maintenance or service personnel can be exposed to an unexpected biohazard should proper containment precautions not be taken. If the affected IABP is not evaluated before use with a new patient, that new patient may be exposed to cross-contaminated blood, which also exposes them to the potential for life-long bacterial and viral diseases such as hepatitis B, hepatitis C, and HIV. Datascope/Getinge has reported 134 complaints about this issue, including 12 device shutdowns and 5 adverse events (4 serious injuries and 1 death)."

## FDA MDR 2024
Base | Device name | Device use | Reason for recall | Exclude/Include - Lara | Exclude/Include - Anninka | (Dis)agreement
FDA MD Recall | High Frequency Transport Phasitron Breathing Circuit Kit (A50605-D) | "The High Frequency Transport Phasitron Breathing Circuit Kit is used for continuous ventilation of patients. The kit is intended to be used with Percussionaire high frequency percussive ventilation (HFPV) systems for either hospital or prehospital use where emergency care is being provided. The kit can also be used for transport within a hospital or outside a hospital. " | "Percussionaire is recalling their High Frequency Transport Phasitron Breathing Circuit Kits due to a product defect that causes over-pressurization. There are two parts of the Phasitron kit that must be pressed together during manufacturing. If one part is angled incorrectly and pressed, the part can become deformed. This causes the PEEP (positive end expiratory pressure) valve to be stuck in the closed position. When the valve is stuck, there is no way to manually release the valve. The use of the affected Phasitron kits may cause serious adverse health consequences, including lung injury, low blood pressure, collapsed lung, cardiac arrest, and death. There have been three complaints regarding this device issue, and one injury. There have been no reports of death."
FDA MD Recall | Atrium Express Dry Suction Dry Seal Chest Drain | "The Express chest drains (models 4000-100N and 4050-100N) are disposable systems used to: Remove air and/or fluid from the chest cavity or mediastinum Help the lungs expand and improve breathing. The Express Chest Drain (Model 4050-100N) facilitates collection of blood from the patient’s pleural cavity or mediastinal area for reinfusion after surgery or to manage blood loss from trauma. Every Express chest drain comes with a pre-packaged 30mL sterile water syringe that is intended to be used to fill the air leak monitor chamber, allowing for the detection of air leaks during or after the initial device set-up, if desired." | "Maquet Cardiovascular, LLC is recalling Atrium Express Dry Suction Dry Seal Chest Drain in direct response to Nurse Assist LLC’s, November 6, 2023, recall of 0.9% sodium chloride irrigation USP and sterile water for irrigation USP, over sterility concerns. If an Express chest drain is knocked over (device not kept in upright position), the water in the air leak monitor chamber could migrate from the air leak monitor chamber to the drainage fluid collection chamber. The patient could potentially be exposed to an infectious pathogen. The use of the affected syringe may cause serious adverse health consequences, including local infection, abscess empyema, mediastinitis, pyothorax, sepsis and death. There have been no reports of injury or death."
FDA MD Recall | ExcelsiusGPS | "ExcelsiusGPS is a surgical guidance system used in spine surgery, combining robotic technology and navigation. It's designed to assist surgeons in accurately placing implants during spinal procedures. The ExcelsiusGPS Flat Panel Fluoroscopy Fixture is a specific component associated with the ExcelsiusGPS surgical guidance system used for inputting medical images during spinal procedures." | "Globus Medical, Inc. is recalling, by removal, the ExcelsiusGPS Flat Panel Fluoroscopy Registration Fixture, a fixture utilized with their ExcelsiusGPS, a stereotaxic navigation device. This recall is being conducted because a calibration process error associated with the ExcelsiusGPS Flat Panel Fluoroscopy Fixture may cause a loss of navigation integrity, which may result in device misplacement and patient harm. When this fixture is used with ExcelsiusGPS to input medical images, there's a chance that the patient's image doesn't align correctly with the computer-assisted design (CAD) model, causing implants or instruments to be misplaced. The root cause is associated with a calibration algorithm error that may affect the accuracy of implant placement. This calibration error was not detected prior to distribution. This misplacement could lead to various complications such as tissue damage, dural tear, cerebrospinal fluid leak, nerve irritation, loss of sensation, paralysis, vascular injury, visceral injury, and pedicle fracture. There have been 2 reported injuries. There have been no reports of death."
FDA MD Recall | Regard Operative Lap P&S Surgical Kit | "The Regard Operative Lap P&S Surgical Kit is a single-use, surgical kit, containing surgical tools and materials for various surgical procedures in different medical areas." | "ROi CPS, LLC is recalling their Regard Operative Lap P&S Surgical Kit due to the possibility they may not be sterile. The surgical kits were built with a recalled item from a separate Class I recall from Nurse Assist. Nurse Assist, LLC recalled their 0.9% sodium chloride irrigation USP and sterile water for irrigation USP due to the potential for a lack of sterility. In populations most at risk, such as patients who have a weakened immune system, there is a possibility that the use of the affected product could potentially result in severe or life-threatening adverse events. There have been zero reports of injury or death."
FDA MD Recall | Ivenix Large Volume Pump | "The Ivenix Large Volume Pump (LVP) is one of three primary components of the Ivenix Infusion System. It's a pump that uses air pressure to precisely control the flow of fluids to the patient. The Ivenix Infusion System is used in hospitals and outpatient centers to carefully give fluids to patients, through different routes like IVs, in arteries, in the spine, or under the skin. It is intended for adults, pediatrics, and newborn babies. The Ivenix Large Volume Pump is compatible only with specific sterile, single-use, disposable administration sets, including: Primary dual inlet, single outlet, 1 Y-site Primary single inlet, single outlet Blood Set, dual inlet, single outlet, 1 Y-site, for administration of red blood cells, platelets and plasma Microbore single inlet, single outlet, for administration of epidural fluids Microbore dual inlet, single outlet, 1 Y-site Primary dual inlet, single outlet, dual Y-site Microbore dual inlet, single outlet, dual Y-site Microbore single inlet, single outlet" | "Fresenius Kabi is recalling the Ivenix Large Volume Pump (LVP) because some units have mechanical issues with the Fluid Valve Pins, located inside the pump's internal housing, causing the pins to not move properly. Instead of going where they should, they impact the side of a sensor. When this happens, the system detects the failure and sets off an alarm, stopping the infusion if it is in progress and preventing the pump from working. The failure may also be detected during the LVP setup, causing a delay in starting treatment. The use of affected product could lead to underdose, interruption in therapy, or delay of therapy, which could lead to serious harm or death. There have been no reports of injuries or death."
FDA MD Recall | AirFit and AirTouch masks | "The AirFit and AirTouch masks are non-continuous ventilatory devices, intended to be used by patients weighing more than 66 lbs. who have been prescribed non-invasive positive airway pressure (PAP) therapy such as CPAP or bi-level therapy. The masks are meant for reuse by one person at home or by multiple people in hospitals. The devices have magnets on the lower headgear straps and frame connections of CPAP masks. These magnets are there to make wearing the mask more comfortable." | "ResMed Ltd. is recalling all their Continuous Positive Airway Pressure (CPAP) masks with magnets due to possible magnetic interference with certain medical devices. Under certain circumstances when a magnet is in close proximity (less than 2 inches) to certain medical implants and devices, it might disrupt their function or position, possibly causing serious harm or death. While the existing label advises keeping magnets 2 inches away from affected medical devices, it doesn't list all the specific ones that could be affected by the masks' magnets. ResMed is recalling these masks to update the labels, add more warnings and information to guide patients and health care professionals on safe usage when using masks with magnets. This recall and label update came after a review of potential risks to medical implants due to magnetic interference. The use of affected masks may cause serious adverse health consequences and death. There have been six reported injuries. There have been no reports of death."
FDA MD Recall | Tracheostomy Care Kit, Dressing Change Kit, and Tracheostomy Care Set | "Busse tracheostomy care trays and kits come sterile packaged and include commonly used components for the removal, cleaning and replacement of the trachea (windpipe) tube along with a dressing change, ensuring convenience and efficiency during medical procedures." | "Busse Hospital Disposables is recalling their Tracheostomy Care Tray, Dressing Change Tray, and Tracheostomy Care Set in direct response to Nurse Assist LLC’s, November 6, 2023, recall of 0.9% sodium chloride irrigation USP and sterile water for irrigation USP, over sterility concerns. This is the only saline used in the Busse trays and kits. The recall doesn't affect the integrity of the other kit components. The use of affected trays or kits may cause serious adverse health consequences, including blood stream infections, urinary tract infections or respiratory infections and in the worst-case scenario sepsis and death. People who have certain health problems like weak immune systems, who are otherwise susceptible to infection or chronic lung diseases may be more at risk. There have been no reports of injury or death."
FDA MD Recall | Megadyne Mega Soft Universal, Megadyne Mega Soft Universal Dual, Megadyne Mega Soft Universal Plus & Megadyne Mega Soft Universal Plus Dual | "The Mega Soft Universal Patient Return Electrode is a medical device for surgeries involving use of electrosurgical instruments. It's designed to ensure the safety of the patient by providing a safe return path for the electrical current used in these procedures. The electrode's purpose is to minimize the risk of electrical burns by ensuring the current flows safely through the patient's body and back to the electrosurgical unit, preventing any concentration of electrical energy in one area of the body. It's an important safety measure in procedures involving electrosurgery to protect the patient from electrical injury." | "Megadyne Medical Product, Inc. is recalling family of Mega Soft Universal Patient Return Electrodes to update the Instructions for Use and product labeling to restrict use of the device to patients 12 years and older. Megadyne has received reports of patient burn injuries, including third-degree burns requiring intervention. Burn injuries may lead to extended hospital stay, scarring, and additional surgeries in both pediatric and adult patients. Severe burns could lead to potentially long-lasting impacts on patients especially under the age of 12 years. Users should continue to follow the current Mega Soft Instructions for Use except for this new limitation in population of intended use. There have been 99 reported injuries. There have been no reports of death."
FDA MD Recall | Omnipod 5 App (on compatible Android smartphones) | "The Omnipod 5 App, available for compatible Android smartphones, is part of the Omnipod 5 System which is an automated insulin delivery system designed to administer insulin to people 2 years of age or older with type 1 diabetes. The Omnipod 5 System includes the Omnipod 5 ACE Pump, the Omnipod 5 App, and a compatible integrated continuous glucose monitor (the Dexcom G6). The Omnipod 5 App uses "SmartAdjust" technology to help adjust insulin delivery based on real-time and future blood sugar levels. The SmartBolus Calculator computes recommended bolus doses. It considers user-inputted carbohydrates, recent sensor glucose (or fingerstick reading), glucose change rate (if available), insulin on board (IOB), and customizable correction factor, insulin-to-carbohydrate ratio, and target glucose value. The recalled Omnipod 5 App is an Android-based software that is provided either on a locked-down Controller or is available to download through Google Play on compatible Android smartphones. This app controls the Omnipod 5 Automated Insulin Delivery System. It is used to activate/deactivate Pods, display alerts/alarms, and send insulin delivery commands for execution to the Pod." | "Insulet Corporation is recalling the Omnipod 5 App for compatible Android smartphones due to a software error that occurs when the user enters a bolus amount less than 1 unit without putting a leading zero before the decimal point. Due to this error, when entering a bolus dose of insulin less than 1 unit, the SmartBolus calculator does not register the decimal point if it is the first character entered. For example, if a user enters ".2" intending to request a 0.2 units insulin bolus, the app would not register the decimal point and would display 2 units as the requested insulin bolus instead. Or if a user enters ".20” units, the app could display 20 units as the requested insulin bolus instead. This failure to recognize the decimal point when entered without a leading zero may lead to the app giving too much insulin, anywhere from 10 times to up to 100 times the intended amount, if the user does not recognize the error on the calculator screen or the confirmation screen before confirming the dose. On November 30, 2023 Insulet Corporation sent all affected customers and health care providers an URGENT Medical Device Correction notice. The letter requested customers to follow instructions when changing a suggested bolus and always visually confirm the bolus amount is critical as the system will always deliver this amount. The firm released a software update v1.2.4 on December 7, 2023 to address and correct the issue. Users should have received a message that an update to the Omnipod 5 App was required and that following December 20, 2023, the application would not be available until they performed the update. The use of affected product without the software update may cause serious adverse health consequences, including severe hypoglycemia and death. There have been 2 reported injuries. There have been no reports of death."